## Task function
Percentage difference from sinc
|  |  | Pixel reduction factor: | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12 | Pixel reduction factor: | 1 | 2 | 3 | 4 | 5 | 6 | 7 | 8 | 9 | 10 | 11 | 12
|  |  | Centre using template match is: 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | [13 empty] | Centre using template match is: 25.6 | 25.6 | Centre using template match is: 25.6 | 25.6 | Centre using template match is: 25.6 | 25.6
|  |  | Final CoM: | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | 25.6 | [13 empty] | Final CoM: | 25.6 | Final CoM: | 25.6 | Final CoM: | 25.6
|  |  | CNR is: ∞ | [25 empty] | CNR is: ∞ |  | CNR is: ∞ |  | CNR is: ∞
|  |  | Contrast is: 1.000 | [25 empty] | Contrast is: 1.000 |  | Contrast is: 1.000 |  | Contrast is: 1.000
|  |  | Noise is: 0.000 | [25 empty] | Noise is: 0.000 |  | Noise is: 0.000 |  | Noise is: 0.000
|  |  | Background is: 0.000 | [25 empty] | Background is: 0.000 |  | Background is: 0.000 |  | Background is: 0.000
|  |  | Signal is: 1.000 | [25 empty] | Signal is: 1.000 |  | Signal is: 1.000 |  | Signal is: 1.000
Pixel size: | 0.1 | mm | MTF 50 and 10 are: 2.101 | 3.554 | 6.529 | 8.165 | 8.981 | 9.551 | 9.779 | 9.788 | 9.553 | 9.779 | 9.561 | 9.711 | 9.693 | [13 empty] | MTF 50 and 10 are: 3.721 | 6.529 | MTF 50 and 10 are: 5.668 | 9.551 | MTF 50 and 10 are: 6.024 | 9.561
|  |  | Freq (mm^-1) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | TTF (Simple rebin) | [13 empty] | Freq (mm^-1) | TTF (Simple rebin) | Freq (mm^-1) | TTF (Simple rebin) | Freq (mm^-1) | TTF (Simple rebin)
Frequency (mm-1) | "x" | sinc = sin(x) / x |  | Plugin (pixel reduction factor of 1) | Plugin (pixel reduction factor of 2) | Plugin (pixel reduction factor of 3) | Plugin (pixel reduction factor of 4) | Plugin (pixel reduction factor of 5) | Plugin (pixel reduction factor of 6) | Plugin (pixel reduction factor of 7) | Plugin (pixel reduction factor of 8) | Plugin (pixel reduction factor of 9) | Plugin (pixel reduction factor of 10) | Plugin (pixel reduction factor of 11) | Plugin (pixel reduction factor of 12) | [13 empty] | 0 | 1 | 0 | 1 | 0 | 1
0 | 0 | 1 | 0 | [13 empty] | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0 | 0.05 | 1 | 0.05 | 1 | 0.05 | 1 | bar: 1 to 12
0.05 | 0.015708 | 0.99996 | 0.05 | [13 empty] | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.0041125 | 0.1 | 0.999 | 0.1 | 1 | 0.1 | 1 | bar: 1 to 12
0.1 | 0.031416 | 0.99984 | 0.1 | 0.998 | 0.999 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | -0.18358 | -0.083565 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.016451 | 0.15 | 0.999 | 0.15 | 1 | 0.15 | 1
0.15 | 0.047124 | 0.99963 | 0.15 | 0.997 | 0.999 | 0.999 | 0.999 | 1 | 1 | 1 | 1 | 1 | 1 | 1 | 1 |  | -0.26309 | -0.063016 | -0.063016 | -0.063016 | 0.037021 | 0.037021 | 0.037021 | 0.037021 | 0.037021 | 0.037021 | 0.037021 | 0.037021 | 0.2 | 0.998 | 0.2 | 0.999 | 0.2 | 0.999
0.2 | 0.062832 | 0.99934 | 0.2 | 0.994 | 0.998 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 |  | -0.53457 | -0.1343 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | -0.034238 | 0.25 | 0.997 | 0.25 | 0.999 | 0.25 | 0.999
0.25 | 0.07854 | 0.99897 | 0.25 | 0.991 | 0.997 | 0.998 | 0.998 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 | 0.999 |  | -0.79804 | -0.19743 | -0.097323 | -0.097323 | 0.0027795 | 0.0027795 | 0.0027795 | 0.0027795 | 0.0027795 | 0.0027795 | 0.0027795 | 0.0027795 | 0.3 | 0.996 | 0.3 | 0.998 | 0.3 | 0.998
0.3 | 0.094248 | 0.99852 | 0.3 | 0.987 | 0.996 | 0.997 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.999 |  | -1.1537 | -0.2524 | -0.15225 | -0.052099 | -0.052099 | -0.052099 | -0.052099 | -0.052099 | -0.052099 | -0.052099 | -0.052099 | 0.048049 | 0.35 | 0.994 | 0.35 | 0.997 | 0.35 | 0.998
0.35 | 0.10996 | 0.99799 | 0.35 | 0.982 | 0.994 | 0.996 | 0.997 | 0.997 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 | 0.998 |  | -1.6018 | -0.39942 | -0.19902 | -0.098816 | -0.098816 | 0.0013854 | 0.0013854 | 0.0013854 | 0.0013854 | 0.0013854 | 0.0013854 | 0.0013854 | 0.4 | 0.992 | 0.4 | 0.997 | 0.4 | 0.997
0.4 | 0.12566 | 0.99737 | 0.4 | 0.977 | 0.992 | 0.995 | 0.996 | 0.997 | 0.997 | 0.997 | 0.997 | 0.997 | 0.997 | 0.997 | 0.997 |  | -2.0424 | -0.53843 | -0.23764 | -0.13738 | -0.037116 | -0.037116 | -0.037116 | -0.037116 | -0.037116 | -0.037116 | -0.037116 | -0.037116 | 0.45 | 0.99 | 0.45 | 0.996 | 0.45 | 0.996
0.45 | 0.14137 | 0.99667 | 0.45 | 0.97 | 0.99 | 0.994 | 0.995 | 0.996 | 0.996 | 0.996 | 0.996 | 0.997 | 0.996 | 0.997 | 0.997 |  | -2.6761 | -0.66946 | -0.26813 | -0.16779 | -0.067458 | -0.067458 | -0.067458 | -0.067458 | 0.032876 | -0.067458 | 0.032876 | 0.032876 | 0.5 | 0.988 | 0.5 | 0.995 | 0.5 | 0.996
0.5 | 0.15708 | 0.99589 | 0.5 | 0.964 | 0.988 | 0.993 | 0.994 | 0.995 | 0.995 | 0.996 | 0.995 | 0.996 | 0.996 | 0.996 | 0.996 |  | -3.2024 | -0.79253 | -0.29047 | -0.19005 | -0.089642 | -0.089642 | 0.010771 | -0.089642 | 0.010771 | 0.010771 | 0.010771 | 0.010771 | 0.55 | 0.985 | 0.55 | 0.994 | 0.55 | 0.995
0.55 | 0.17279 | 0.99503 | 0.55 | 0.956 | 0.985 | 0.991 | 0.993 | 0.994 | 0.994 | 0.995 | 0.995 | 0.995 | 0.995 | 0.995 | 0.995 |  | -3.9226 | -1.0082 | -0.40516 | -0.20416 | -0.10366 | -0.10366 | -0.0031654 | -0.0031654 | -0.0031654 | -0.0031654 | -0.0031654 | -0.0031654 | 0.6 | 0.983 | 0.6 | 0.993 | 0.6 | 0.994
0.6 | 0.1885 | 0.99409 | 0.6 | 0.948 | 0.983 | 0.989 | 0.991 | 0.993 | 0.993 | 0.994 | 0.994 | 0.994 | 0.994 | 0.994 | 0.994 |  | -4.6363 | -1.1155 | -0.5119 | -0.31071 | -0.10952 | -0.10952 | -0.0089276 | -0.0089276 | -0.0089276 | -0.0089276 | -0.0089276 | -0.0089276 | 0.65 | 0.98 | 0.65 | 0.991 | 0.65 | 0.993
0.65 | 0.2042 | 0.99306 | 0.65 | 0.939 | 0.98 | 0.988 | 0.99 | 0.991 | 0.992 | 0.992 | 0.992 | 0.993 | 0.993 | 0.993 | 0.993 |  | -5.4442 | -1.3156 | -0.51 | -0.3086 | -0.2079 | -0.10721 | -0.10721 | -0.10721 | -0.0065081 | -0.0065081 | -0.0065081 | -0.0065081 | 0.7 | 0.976 | 0.7 | 0.99 | 0.7 | 0.992
0.7 | 0.21991 | 0.99196 | 0.7 | 0.93 | 0.976 | 0.986 | 0.988 | 0.99 | 0.991 | 0.991 | 0.991 | 0.992 | 0.992 | 0.992 | 0.992 |  | -6.2462 | -1.6089 | -0.60076 | -0.39914 | -0.19752 | -0.096707 | -0.096707 | -0.096707 | 0.0041039 | 0.0041039 | 0.0041039 | 0.0041039 | 0.75 | 0.973 | 0.75 | 0.989 | 0.75 | 0.99
0.75 | 0.23562 | 0.99077 | 0.75 | 0.92 | 0.973 | 0.984 | 0.987 | 0.989 | 0.99 | 0.99 | 0.99 | 0.99 | 0.99 | 0.99 | 0.991 |  | -7.1432 | -1.7938 | -0.6836 | -0.3808 | -0.17894 | -0.078009 | -0.078009 | -0.078009 | -0.078009 | -0.078009 | -0.078009 | 0.022922 | 0.8 | 0.969 | 0.8 | 0.987 | 0.8 | 0.989
0.8 | 0.25133 | 0.98951 | 0.8 | 0.909 | 0.969 | 0.981 | 0.985 | 0.987 | 0.988 | 0.989 | 0.988 | 0.989 | 0.989 | 0.989 | 0.989 |  | -8.1359 | -2.0723 | -0.85958 | -0.45534 | -0.25322 | -0.15216 | -0.051098 | -0.15216 | -0.051098 | -0.051098 | -0.051098 | -0.051098 | 0.85 | 0.965 | 0.85 | 0.985 | 0.85 | 0.988
0.85 | 0.26704 | 0.98816 | 0.85 | 0.898 | 0.965 | 0.979 | 0.983 | 0.985 | 0.987 | 0.987 | 0.987 | 0.988 | 0.988 | 0.988 | 0.988 |  | -9.1238 | -2.3435 | -0.92674 | -0.52195 | -0.31955 | -0.11715 | -0.11715 | -0.11715 | -0.015954 | -0.015954 | -0.015954 | -0.015954 | 0.9 | 0.961 | 0.9 | 0.984 | 0.9 | 0.986
0.9 | 0.28274 | 0.98673 | 0.9 | 0.886 | 0.961 | 0.976 | 0.981 | 0.984 | 0.985 | 0.985 | 0.985 | 0.986 | 0.986 | 0.986 | 0.987 |  | -10.2084 | -2.6075 | -1.0873 | -0.58062 | -0.27659 | -0.17524 | -0.17524 | -0.17524 | -0.0739 | -0.0739 | -0.0739 | 0.027445 | 0.95 | 0.957 | 0.95 | 0.982 | 0.95 | 0.984
0.95 | 0.29845 | 0.98522 | 0.95 | 0.874 | 0.957 | 0.974 | 0.979 | 0.982 | 0.983 | 0.984 | 0.984 | 0.985 | 0.984 | 0.985 | 0.985 |  | -11.2889 | -2.8644 | -1.1389 | -0.63138 | -0.32688 | -0.22538 | -0.12388 | -0.12388 | -0.022375 | -0.12388 | -0.022375 | -0.022375 | 1 | 0.953 | 1 | 0.98 | 1 | 0.983
1 | 0.31416 | 0.98363 | 1 | 0.861 | 0.953 | 0.971 | 0.976 | 0.98 | 0.981 | 0.982 | 0.982 | 0.983 | 0.983 | 0.983 | 0.983 |  | -12.4672 | -3.1141 | -1.2842 | -0.77586 | -0.36921 | -0.26754 | -0.16588 | -0.16588 | -0.064215 | -0.064215 | -0.064215 | -0.064215 | 1.05 | 0.948 | 1.05 | 0.978 | 1.05 | 0.981
1.05 | 0.32987 | 0.98196 | 1.05 | 0.848 | 0.948 | 0.968 | 0.974 | 0.978 | 0.98 | 0.98 | 0.98 | 0.981 | 0.981 | 0.981 | 0.982 |  | -13.6424 | -3.4587 | -1.4219 | -0.81093 | -0.40358 | -0.19991 | -0.19991 | -0.19991 | -0.09807 | -0.09807 | -0.09807 | 0.0037665 | 1.1 | 0.943 | 1.1 | 0.976 | 1.1 | 0.979
1.1 | 0.34558 | 0.98021 | 1.1 | 0.834 | 0.943 | 0.965 | 0.971 | 0.976 | 0.978 | 0.978 | 0.978 | 0.98 | 0.979 | 0.98 | 0.98 |  | -14.9166 | -3.7966 | -1.5522 | -0.94008 | -0.42999 | -0.22595 | -0.22595 | -0.22595 | -0.021914 | -0.12393 | -0.021914 | -0.021914 | 1.15 | 0.938 | 1.15 | 0.973 | 1.15 | 0.977
1.15 | 0.36128 | 0.97839 | 1.15 | 0.82 | 0.938 | 0.962 | 0.969 | 0.973 | 0.976 | 0.976 | 0.976 | 0.978 | 0.977 | 0.978 | 0.978 |  | -16.1886 | -4.1279 | -1.6749 | -0.95946 | -0.55063 | -0.244 | -0.244 | -0.244 | -0.039584 | -0.14179 | -0.039584 | -0.039584 | 1.2 | 0.932 | 1.2 | 0.971 | 1.2 | 0.975
1.2 | 0.37699 | 0.97648 | 1.2 | 0.805 | 0.932 | 0.958 | 0.966 | 0.971 | 0.973 | 0.974 | 0.974 | 0.976 | 0.975 | 0.976 | 0.976 |  | -17.5611 | -4.5552 | -1.8926 | -1.0733 | -0.56127 | -0.35645 | -0.25405 | -0.25405 | -0.049228 | -0.15164 | -0.049228 | -0.049228 | 1.25 | 0.927 | 1.25 | 0.969 | 1.25 | 0.973
1.25 | 0.3927 | 0.9745 | 1.25 | 0.79 | 0.927 | 0.955 | 0.963 | 0.969 | 0.971 | 0.972 | 0.972 | 0.974 | 0.973 | 0.974 | 0.974 |  | -18.9324 | -4.8738 | -2.0006 | -1.1796 | -0.56392 | -0.35868 | -0.25607 | -0.25607 | -0.050832 | -0.15345 | -0.050832 | -0.050832 | 1.3 | 0.921 | 1.3 | 0.966 | 1.3 | 0.971
1.3 | 0.40841 | 0.97243 | 1.3 | 0.775 | 0.921 | 0.951 | 0.96 | 0.966 | 0.969 | 0.97 | 0.97 | 0.971 | 0.971 | 0.972 | 0.972 |  | -20.3029 | -5.289 | -2.2039 | -1.2784 | -0.66139 | -0.35288 | -0.25005 | -0.25005 | -0.14721 | -0.14721 | -0.044377 | -0.044377 | 1.35 | 0.915 | 1.35 | 0.963 | 1.35 | 0.969
1.35 | 0.42412 | 0.97029 | 1.35 | 0.759 | 0.915 | 0.947 | 0.957 | 0.963 | 0.966 | 0.968 | 0.968 | 0.969 | 0.969 | 0.969 | 0.97 |  | -21.7759 | -5.6983 | -2.4003 | -1.3696 | -0.75128 | -0.44209 | -0.23597 | -0.23597 | -0.1329 | -0.1329 | -0.1329 | -0.029841 | 1.4 | 0.909 | 1.4 | 0.961 | 1.4 | 0.966
1.4 | 0.43982 | 0.96807 | 1.4 | 0.743 | 0.909 | 0.944 | 0.954 | 0.961 | 0.964 | 0.965 | 0.965 | 0.967 | 0.966 | 0.967 | 0.968 |  | -23.2493 | -6.1018 | -2.4864 | -1.4534 | -0.73029 | -0.42039 | -0.31709 | -0.31709 | -0.1105 | -0.2138 | -0.1105 | -0.0071998 | 1.45 | 0.903 | 1.45 | 0.958 | 1.45 | 0.964
1.45 | 0.45553 | 0.96577 | 1.45 | 0.727 | 0.903 | 0.94 | 0.951 | 0.958 | 0.961 | 0.963 | 0.963 | 0.965 | 0.964 | 0.965 | 0.965 |  | -24.7235 | -6.4997 | -2.6686 | -1.5296 | -0.80478 | -0.49415 | -0.28706 | -0.28706 | -0.07997 | -0.18351 | -0.07997 | -0.07997 | 1.5 | 0.896 | 1.5 | 0.955 | 1.5 | 0.961
1.5 | 0.47124 | 0.9634 | 1.5 | 0.71 | 0.896 | 0.935 | 0.947 | 0.955 | 0.959 | 0.96 | 0.96 | 0.962 | 0.961 | 0.962 | 0.963 |  | -26.3025 | -6.9958 | -2.9477 | -1.7021 | -0.87168 | -0.45648 | -0.35269 | -0.35269 | -0.14509 | -0.24889 | -0.14509 | -0.041287 | 1.55 | 0.889 | 1.55 | 0.952 | 1.55 | 0.959
1.55 | 0.48695 | 0.96095 | 1.55 | 0.693 | 0.889 | 0.931 | 0.944 | 0.952 | 0.956 | 0.957 | 0.957 | 0.96 | 0.959 | 0.96 | 0.961 |  | -27.8836 | -7.487 | -3.1163 | -1.7635 | -0.93099 | -0.51474 | -0.41067 | -0.41067 | -0.098482 | -0.20255 | -0.098482 | 0.0055819 | 1.6 | 0.883 | 1.6 | 0.949 | 1.6 | 0.956
1.6 | 0.50265 | 0.95842 | 1.6 | 0.676 | 0.883 | 0.927 | 0.94 | 0.949 | 0.953 | 0.955 | 0.955 | 0.957 | 0.956 | 0.957 | 0.958 |  | -29.4671 | -7.8691 | -3.2782 | -1.9218 | -0.98271 | -0.56536 | -0.35668 | -0.35668 | -0.148 | -0.25234 | -0.148 | -0.043664 | 1.65 | 0.876 | 1.65 | 0.946 | 1.65 | 0.954
1.65 | 0.51836 | 0.95581 | 1.65 | 0.659 | 0.876 | 0.922 | 0.936 | 0.946 | 0.95 | 0.952 | 0.952 | 0.954 | 0.954 | 0.955 | 0.955 |  | -31.0536 | -8.3504 | -3.5378 | -2.073 | -1.0268 | -0.60833 | -0.39908 | -0.39908 | -0.18984 | -0.18984 | -0.085215 | -0.085215 | 1.7 | 0.868 | 1.7 | 0.942 | 1.7 | 0.951
1.7 | 0.53407 | 0.95313 | 1.7 | 0.641 | 0.868 | 0.918 | 0.933 | 0.942 | 0.947 | 0.949 | 0.949 | 0.952 | 0.951 | 0.952 | 0.953 |  | -32.7482 | -8.9321 | -3.6862 | -2.1125 | -1.1682 | -0.64364 | -0.43381 | -0.43381 | -0.11906 | -0.22398 | -0.11906 | -0.014142 | 1.75 | 0.861 | 1.75 | 0.939 | 1.75 | 0.948
1.75 | 0.54978 | 0.95038 | 1.75 | 0.624 | 0.861 | 0.913 | 0.929 | 0.939 | 0.944 | 0.946 | 0.946 | 0.949 | 0.948 | 0.949 | 0.95 |  | -34.342 | -9.4046 | -3.9331 | -2.2496 | -1.1974 | -0.67129 | -0.46084 | -0.46084 | -0.14518 | -0.2504 | -0.14518 | -0.039959 | 1.8 | 0.854 | 1.8 | 0.936 | 1.8 | 0.945
1.8 | 0.56549 | 0.94755 | 1.8 | 0.606 | 0.854 | 0.908 | 0.925 | 0.936 | 0.941 | 0.943 | 0.943 | 0.946 | 0.945 | 0.946 | 0.947 |  | -36.0456 | -9.8728 | -4.1739 | -2.3798 | -1.2189 | -0.69124 | -0.48017 | -0.48017 | -0.16356 | -0.2691 | -0.16356 | -0.058025 | 1.85 | 0.846 | 1.85 | 0.932 | 1.85 | 0.942
1.85 | 0.58119 | 0.94465 | 1.85 | 0.589 | 0.846 | 0.903 | 0.921 | 0.932 | 0.938 | 0.94 | 0.94 | 0.943 | 0.942 | 0.943 | 0.944 |  | -37.6486 | -10.4426 | -4.4086 | -2.5031 | -1.3386 | -0.70348 | -0.49176 | -0.49176 | -0.17418 | -0.28004 | -0.17418 | -0.068317 | 1.9 | 0.838 | 1.9 | 0.929 | 1.9 | 0.939
1.9 | 0.5969 | 0.94167 | 1.9 | 0.571 | 0.838 | 0.898 | 0.917 | 0.929 | 0.935 | 0.937 | 0.936 | 0.94 | 0.939 | 0.94 | 0.941 |  | -39.3628 | -11.0089 | -4.6372 | -2.6195 | -1.3451 | -0.70798 | -0.49559 | -0.60179 | -0.17701 | -0.2832 | -0.17701 | -0.070813 | 1.95 | 0.83 | 1.95 | 0.925 | 1.95 | 0.936
1.95 | 0.61261 | 0.93861 | 1.95 | 0.553 | 0.83 | 0.893 | 0.912 | 0.925 | 0.931 | 0.933 | 0.933 | 0.937 | 0.936 | 0.937 | 0.938 |  | -41.0834 | -11.5718 | -4.8598 | -2.8355 | -1.4505 | -0.81126 | -0.59818 | -0.59818 | -0.17202 | -0.27856 | -0.17202 | -0.065485 | 2 | 0.822 | 2 | 0.921 | 2 | 0.932
2 | 0.62832 | 0.93549 | 2 | 0.536 | 0.822 | 0.888 | 0.908 | 0.921 | 0.928 | 0.93 | 0.93 | 0.933 | 0.932 | 0.934 | 0.935 |  | -42.7038 | -12.1315 | -5.0764 | -2.9385 | -1.5488 | -0.80057 | -0.58678 | -0.58678 | -0.26609 | -0.37299 | -0.1592 | -0.052302 | 2.05 | 0.814 | 2.05 | 0.917 | 2.05 | 0.929
2.05 | 0.64403 | 0.93229 | 2.05 | 0.518 | 0.814 | 0.882 | 0.903 | 0.917 | 0.924 | 0.927 | 0.926 | 0.93 | 0.929 | 0.931 | 0.932 |  | -44.438 | -12.6882 | -5.3944 | -3.1418 | -1.6402 | -0.88933 | -0.56755 | -0.67481 | -0.24576 | -0.35302 | -0.1385 | -0.031234 | 2.1 | 0.806 | 2.1 | 0.913 | 2.1 | 0.926
2.1 | 0.65973 | 0.92902 | 2.1 | 0.5 | 0.806 | 0.877 | 0.899 | 0.913 | 0.92 | 0.923 | 0.923 | 0.927 | 0.926 | 0.927 | 0.929 |  | -46.1799 | -13.242 | -5.5995 | -3.2314 | -1.7245 | -0.971 | -0.64808 | -0.64808 | -0.21752 | -0.32516 | -0.21752 | -0.0022422 | 2.15 | 0.797 | 2.15 | 0.909 | 2.15 | 0.922
2.15 | 0.67544 | 0.92568 | 2.15 | 0.483 | 0.797 | 0.871 | 0.894 | 0.909 | 0.917 | 0.92 | 0.919 | 0.923 | 0.922 | 0.924 | 0.925 |  | -47.8221 | -13.901 | -5.9069 | -3.4222 | -1.8018 | -0.93755 | -0.61346 | -0.72149 | -0.28938 | -0.3974 | -0.18135 | -0.073318 | 2.2 | 0.789 | 2.2 | 0.905 | 2.2 | 0.919
2.2 | 0.69115 | 0.92227 | 2.2 | 0.465 | 0.789 | 0.866 | 0.89 | 0.905 | 0.913 | 0.916 | 0.916 | 0.92 | 0.919 | 0.92 | 0.922 |  | -49.5807 | -14.4498 | -6.1008 | -3.4985 | -1.872 | -1.0046 | -0.67933 | -0.67933 | -0.24562 | -0.35405 | -0.24562 | -0.028762 | 2.25 | 0.78 | 2.25 | 0.901 | 2.25 | 0.915
2.25 | 0.70686 | 0.91878 | 2.25 | 0.448 | 0.78 | 0.86 | 0.885 | 0.901 | 0.909 | 0.912 | 0.912 | 0.916 | 0.915 | 0.917 | 0.918 |  | -51.2397 | -15.1049 | -6.3977 | -3.6767 | -1.9353 | -1.0646 | -0.73805 | -0.73805 | -0.30269 | -0.41153 | -0.19385 | -0.085008 | 2.3 | 0.771 | 2.3 | 0.897 | 2.3 | 0.911
2.3 | 0.72257 | 0.91523 | 2.3 | 0.431 | 0.771 | 0.854 | 0.88 | 0.897 | 0.905 | 0.908 | 0.908 | 0.913 | 0.911 | 0.913 | 0.915 |  | -52.9078 | -15.7586 | -6.6898 | -3.8489 | -1.9915 | -1.1174 | -0.78959 | -0.78959 | -0.24328 | -0.4618 | -0.24328 | -0.024753 | 2.35 | 0.762 | 2.35 | 0.892 | 2.35 | 0.907
2.35 | 0.73827 | 0.9116 | 2.35 | 0.414 | 0.762 | 0.848 | 0.875 | 0.892 | 0.901 | 0.904 | 0.904 | 0.909 | 0.907 | 0.909 | 0.911 |  | -54.5855 | -16.4109 | -6.977 | -4.0152 | -2.1503 | -1.163 | -0.83395 | -0.83395 | -0.28546 | -0.50486 | -0.28546 | -0.066069 | 2.4 | 0.754 | 2.4 | 0.888 | 2.4 | 0.904
2.4 | 0.75398 | 0.90791 | 2.4 | 0.397 | 0.754 | 0.842 | 0.87 | 0.888 | 0.897 | 0.901 | 0.9 | 0.905 | 0.904 | 0.906 | 0.908 |  | -56.2731 | -16.952 | -7.2594 | -4.1754 | -2.1928 | -1.2015 | -0.76096 | -0.8711 | -0.32038 | -0.43053 | -0.21024 | 0.010045 | 2.45 | 0.744 | 2.45 | 0.884 | 2.45 | 0.9
2.45 | 0.76969 | 0.90415 | 2.45 | 0.38 | 0.744 | 0.836 | 0.865 | 0.884 | 0.893 | 0.897 | 0.896 | 0.901 | 0.9 | 0.902 | 0.904 |  | -57.9714 | -17.7125 | -7.5371 | -4.3297 | -2.2282 | -1.2328 | -0.79043 | -0.90103 | -0.34802 | -0.45862 | -0.23742 | -0.016217 | 2.5 | 0.735 | 2.5 | 0.879 | 2.5 | 0.896
2.5 | 0.7854 | 0.90032 | 2.5 | 0.364 | 0.735 | 0.83 | 0.859 | 0.879 | 0.889 | 0.892 | 0.892 | 0.898 | 0.896 | 0.898 | 0.9 |  | -59.5698 | -18.362 | -7.8102 | -4.5891 | -2.3676 | -1.2569 | -0.92371 | -0.92371 | -0.25728 | -0.47942 | -0.25728 | -0.035134 | 2.55 | 0.726 | 2.55 | 0.874 | 2.55 | 0.892
2.55 | 0.80111 | 0.89642 | 2.55 | 0.348 | 0.726 | 0.823 | 0.854 | 0.874 | 0.885 | 0.888 | 0.888 | 0.894 | 0.892 | 0.894 | 0.896 |  | -61.1788 | -19.011 | -8.1902 | -4.732 | -2.5009 | -1.2738 | -0.93912 | -0.93912 | -0.26979 | -0.4929 | -0.26979 | -0.046678 | 2.6 | 0.717 | 2.6 | 0.87 | 2.6 | 0.888
2.6 | 0.81681 | 0.89245 | 2.6 | 0.332 | 0.717 | 0.817 | 0.849 | 0.87 | 0.88 | 0.884 | 0.884 | 0.89 | 0.888 | 0.89 | 0.892 |  | -62.7992 | -19.6597 | -8.4546 | -4.869 | -2.5159 | -1.3954 | -0.94722 | -0.94722 | -0.27492 | -0.49902 | -0.27492 | -0.05082 | 2.65 | 0.707 | 2.65 | 0.865 | 2.65 | 0.883
2.65 | 0.83252 | 0.88842 | 2.65 | 0.317 | 0.707 | 0.81 | 0.843 | 0.865 | 0.876 | 0.88 | 0.879 | 0.885 | 0.883 | 0.886 | 0.888 |  | -64.3188 | -20.4207 | -8.8271 | -5.1127 | -2.6364 | -1.3982 | -0.948 | -1.0606 | -0.3852 | -0.61032 | -0.27264 | -0.047525 | 2.7 | 0.698 | 2.7 | 0.86 | 2.7 | 0.879
2.7 | 0.84823 | 0.88433 | 2.7 | 0.302 | 0.698 | 0.804 | 0.838 | 0.86 | 0.871 | 0.875 | 0.875 | 0.881 | 0.879 | 0.882 | 0.884 |  | -65.8497 | -21.0697 | -9.0832 | -5.2385 | -2.7507 | -1.5068 | -1.0545 | -1.0545 | -0.376 | -0.60216 | -0.26292 | -0.036759 | 2.75 | 0.689 | 2.75 | 0.855 | 2.75 | 0.875
2.75 | 0.86394 | 0.88016 | 2.75 | 0.287 | 0.689 | 0.797 | 0.832 | 0.855 | 0.867 | 0.871 | 0.871 | 0.877 | 0.875 | 0.878 | 0.88 |  | -67.3924 | -21.719 | -9.4486 | -5.472 | -2.8589 | -1.4955 | -1.041 | -1.041 | -0.35933 | -0.58656 | -0.24571 | -0.018483 | 2.8 | 0.679 | 2.8 | 0.85 | 2.8 | 0.871
2.8 | 0.87965 | 0.87594 | 2.8 | 0.272 | 0.679 | 0.791 | 0.827 | 0.85 | 0.862 | 0.867 | 0.866 | 0.873 | 0.871 | 0.873 | 0.876 |  | -68.9475 | -22.4829 | -9.6966 | -5.5867 | -2.9609 | -1.5909 | -1.0201 | -1.1343 | -0.33514 | -0.56347 | -0.33514 | 0.0073462 | 2.85 | 0.669 | 2.85 | 0.845 | 2.85 | 0.866
2.85 | 0.89535 | 0.87164 | 2.85 | 0.258 | 0.669 | 0.784 | 0.821 | 0.845 | 0.858 | 0.862 | 0.861 | 0.868 | 0.866 | 0.869 | 0.871 |  | -70.4008 | -23.2485 | -10.0551 | -5.8102 | -3.0568 | -1.5654 | -1.1065 | -1.2212 | -0.41813 | -0.64758 | -0.3034 | -0.073953 | 2.9 | 0.66 | 2.9 | 0.84 | 2.9 | 0.862
2.9 | 0.91106 | 0.86729 | 2.9 | 0.244 | 0.66 | 0.777 | 0.815 | 0.84 | 0.853 | 0.857 | 0.857 | 0.864 | 0.862 | 0.865 | 0.867 |  | -71.8664 | -23.9009 | -10.4106 | -6.0291 | -3.1466 | -1.6477 | -1.1865 | -1.1865 | -0.37936 | -0.60997 | -0.26406 | -0.033458 | 2.95 | 0.65 | 2.95 | 0.835 | 2.95 | 0.857
2.95 | 0.92677 | 0.86287 | 2.95 | 0.231 | 0.65 | 0.77 | 0.809 | 0.835 | 0.848 | 0.853 | 0.852 | 0.86 | 0.857 | 0.86 | 0.863 |  | -73.229 | -24.6703 | -10.7632 | -6.2434 | -3.2303 | -1.7237 | -1.1442 | -1.2601 | -0.33296 | -0.68063 | -0.33296 | 0.014719 | 3 | 0.641 | 3 | 0.83 | 3 | 0.853
3 | 0.94248 | 0.85839 | 3 | 0.218 | 0.641 | 0.763 | 0.803 | 0.83 | 0.843 | 0.848 | 0.847 | 0.855 | 0.853 | 0.856 | 0.858 |  | -74.6037 | -25.3256 | -11.113 | -6.4532 | -3.3078 | -1.7933 | -1.2108 | -1.3273 | -0.39535 | -0.62835 | -0.27886 | -0.045864 | 3.05 | 0.631 | 3.05 | 0.825 | 3.05 | 0.848
3.05 | 0.95819 | 0.85385 | 3.05 | 0.205 | 0.631 | 0.756 | 0.797 | 0.825 | 0.838 | 0.843 | 0.843 | 0.85 | 0.848 | 0.851 | 0.854 |  | -75.9912 | -26.0997 | -11.4602 | -6.6584 | -3.3791 | -1.8566 | -1.2711 | -1.2711 | -0.45124 | -0.68547 | -0.33412 | 0.017224 | 3.1 | 0.621 | 3.1 | 0.819 | 3.1 | 0.843
3.1 | 0.97389 | 0.84925 | 3.1 | 0.193 | 0.621 | 0.749 | 0.791 | 0.819 | 0.833 | 0.839 | 0.838 | 0.846 | 0.843 | 0.846 | 0.849 |  | -77.2741 | -26.8768 | -11.8047 | -6.8591 | -3.5621 | -1.9136 | -1.2071 | -1.3249 | -0.38285 | -0.7361 | -0.38285 | -0.029599 | 3.15 | 0.611 | 3.15 | 0.814 | 3.15 | 0.838
3.15 | 0.9896 | 0.84459 | 3.15 | 0.181 | 0.611 | 0.742 | 0.785 | 0.814 | 0.828 | 0.834 | 0.833 | 0.841 | 0.838 | 0.842 | 0.845 |  | -78.5695 | -27.6572 | -12.1467 | -7.0555 | -3.6218 | -1.9642 | -1.2538 | -1.3722 | -0.42502 | -0.78022 | -0.30662 | 0.048583 | 3.2 | 0.601 | 3.2 | 0.809 | 3.2 | 0.834
3.2 | 1.0053 | 0.83987 | 3.2 | 0.169 | 0.601 | 0.735 | 0.779 | 0.809 | 0.823 | 0.829 | 0.828 | 0.836 | 0.834 | 0.837 | 0.84 |  | -79.8778 | -28.4412 | -12.4863 | -7.2474 | -3.6754 | -2.0085 | -1.2941 | -1.4131 | -0.46061 | -0.69874 | -0.34154 | 0.015655 | 3.25 | 0.592 | 3.25 | 0.803 | 3.25 | 0.829
3.25 | 1.021 | 0.83509 | 3.25 | 0.158 | 0.592 | 0.727 | 0.773 | 0.803 | 0.818 | 0.824 | 0.823 | 0.832 | 0.829 | 0.832 | 0.835 |  | -81.0799 | -29.1093 | -12.9434 | -7.435 | -3.8425 | -2.0463 | -1.3278 | -1.4476 | -0.36985 | -0.7291 | -0.36985 | -0.010609 | 3.3 | 0.582 | 3.3 | 0.798 | 3.3 | 0.824
3.3 | 1.0367 | 0.83025 | 3.3 | 0.148 | 0.582 | 0.72 | 0.766 | 0.798 | 0.813 | 0.819 | 0.818 | 0.827 | 0.824 | 0.827 | 0.831 |  | -82.1741 | -29.9007 | -13.2792 | -7.7387 | -3.8844 | -2.0778 | -1.3551 | -1.4755 | -0.39152 | -0.75286 | -0.39152 | 0.090262 | 3.35 | 0.572 | 3.35 | 0.792 | 3.35 | 0.819
3.35 | 1.0524 | 0.82536 | 3.35 | 0.137 | 0.572 | 0.713 | 0.76 | 0.792 | 0.808 | 0.814 | 0.813 | 0.822 | 0.819 | 0.823 | 0.826 |  | -83.4011 | -30.6965 | -13.613 | -7.9184 | -4.0413 | -2.1028 | -1.3758 | -1.497 | -0.40652 | -0.77 | -0.28536 | 0.078122 | 3.4 | 0.562 | 3.4 | 0.786 | 3.4 | 0.814
3.4 | 1.0681 | 0.8204 | 3.4 | 0.127 | 0.562 | 0.705 | 0.754 | 0.786 | 0.803 | 0.808 | 0.807 | 0.817 | 0.814 | 0.818 | 0.821 |  | -84.5198 | -31.4971 | -14.0666 | -8.094 | -4.1934 | -2.1213 | -1.5118 | -1.6337 | -0.41482 | -0.78049 | -0.29293 | 0.072748 | 3.45 | 0.553 | 3.45 | 0.781 | 3.45 | 0.809
3.45 | 1.0838 | 0.8154 | 3.45 | 0.118 | 0.553 | 0.698 | 0.747 | 0.781 | 0.797 | 0.803 | 0.802 | 0.812 | 0.809 | 0.813 | 0.816 |  | -85.5285 | -32.1801 | -14.3973 | -8.388 | -4.2182 | -2.256 | -1.5201 | -1.6428 | -0.41638 | -0.7843 | -0.29374 | 0.074175 | 3.5 | 0.543 | 3.5 | 0.775 | 3.5 | 0.803
3.5 | 1.0996 | 0.81033 | 3.5 | 0.109 | 0.543 | 0.691 | 0.741 | 0.775 | 0.792 | 0.798 | 0.797 | 0.807 | 0.803 | 0.808 | 0.811 |  | -86.5487 | -32.9904 | -14.7263 | -8.556 | -4.3602 | -2.2623 | -1.5218 | -1.6452 | -0.41118 | -0.90481 | -0.28778 | 0.082441 | 3.55 | 0.533 | 3.55 | 0.769 | 3.55 | 0.798
3.55 | 1.1153 | 0.80521 | 3.55 | 0.101 | 0.533 | 0.683 | 0.734 | 0.769 | 0.787 | 0.793 | 0.792 | 0.802 | 0.798 | 0.802 | 0.806 |  | -87.4568 | -33.8064 | -15.1779 | -8.8441 | -4.4975 | -2.262 | -1.5169 | -1.6411 | -0.39918 | -0.89594 | -0.39918 | 0.097584 | 3.6 | 0.523 | 3.6 | 0.763 | 3.6 | 0.793
3.6 | 1.131 | 0.80004 | 3.6 | 0.092 | 0.523 | 0.675 | 0.728 | 0.763 | 0.781 | 0.787 | 0.786 | 0.797 | 0.793 | 0.797 | 0.801 |  | -88.5006 | -34.6285 | -15.6295 | -9.0049 | -4.6301 | -2.3802 | -1.6303 | -1.7553 | -0.38033 | -0.8803 | -0.38033 | 0.11965 | 3.65 | 0.514 | 3.65 | 0.757 | 3.65 | 0.788
3.65 | 1.1467 | 0.79482 | 3.65 | 0.085 | 0.514 | 0.668 | 0.721 | 0.757 | 0.776 | 0.782 | 0.781 | 0.791 | 0.788 | 0.792 | 0.796 |  | -89.3057 | -35.3311 | -15.9556 | -9.2874 | -4.7581 | -2.3676 | -1.6127 | -1.7385 | -0.4804 | -0.85784 | -0.35458 | 0.14868 | 3.7 | 0.504 | 3.7 | 0.751 | 3.7 | 0.782
3.7 | 1.1624 | 0.78954 | 3.7 | 0.077 | 0.504 | 0.66 | 0.715 | 0.751 | 0.77 | 0.777 | 0.775 | 0.786 | 0.782 | 0.787 | 0.791 |  | -90.2475 | -36.1655 | -16.4072 | -9.4411 | -4.8815 | -2.475 | -1.5885 | -1.8418 | -0.44856 | -0.95518 | -0.3219 | 0.18472 | 3.75 | 0.494 | 3.75 | 0.746 | 3.75 | 0.777
3.75 | 1.1781 | 0.78421 | 3.75 | 0.071 | 0.494 | 0.653 | 0.708 | 0.746 | 0.764 | 0.771 | 0.77 | 0.781 | 0.777 | 0.782 | 0.785 |  | -90.9463 | -37.0069 | -16.7318 | -9.7184 | -4.8728 | -2.5775 | -1.6849 | -1.8124 | -0.40975 | -0.91981 | -0.28223 | 0.10032 | 3.8 | 0.485 | 3.8 | 0.74 | 3.8 | 0.772
3.8 | 1.1938 | 0.77883 | 3.8 | 0.064 | 0.485 | 0.645 | 0.701 | 0.74 | 0.759 | 0.766 | 0.764 | 0.775 | 0.772 | 0.776 | 0.78 |  | -91.7826 | -37.7275 | -17.1839 | -9.9937 | -4.9862 | -2.5467 | -1.6479 | -1.9047 | -0.49232 | -0.87751 | -0.36392 | 0.14967 | 3.85 | 0.475 | 3.85 | 0.733 | 3.85 | 0.766
3.85 | 1.2095 | 0.77341 | 3.85 | 0.058 | 0.475 | 0.637 | 0.695 | 0.733 | 0.753 | 0.76 | 0.759 | 0.77 | 0.766 | 0.771 | 0.775 |  | -92.5007 | -38.5833 | -17.637 | -10.1377 | -5.2244 | -2.6384 | -1.7333 | -1.8626 | -0.44031 | -0.95751 | -0.31102 | 0.20618 | 3.9 | 0.466 | 3.9 | 0.727 | 3.9 | 0.761
3.9 | 1.2252 | 0.76793 | 3.9 | 0.052 | 0.466 | 0.63 | 0.688 | 0.727 | 0.747 | 0.754 | 0.753 | 0.765 | 0.761 | 0.765 | 0.769 |  | -93.2285 | -39.3172 | -17.961 | -10.4082 | -5.3296 | -2.7252 | -1.8136 | -1.9438 | -0.3812 | -0.90208 | -0.3812 | 0.13969 | 3.95 | 0.456 | 3.95 | 0.721 | 3.95 | 0.755
3.95 | 1.2409 | 0.7624 | 3.95 | 0.047 | 0.456 | 0.622 | 0.681 | 0.721 | 0.742 | 0.749 | 0.747 | 0.759 | 0.755 | 0.76 | 0.764 |  | -93.8353 | -40.1889 | -18.4156 | -10.6769 | -5.4303 | -2.6759 | -1.7577 | -2.02 | -0.44607 | -0.97073 | -0.3149 | 0.20976 | 4 | 0.447 | 4 | 0.715 | 4 | 0.749
4 | 1.2566 | 0.75683 | 4 | 0.042 | 0.447 | 0.614 | 0.674 | 0.715 | 0.736 | 0.743 | 0.742 | 0.754 | 0.749 | 0.754 | 0.759 |  | -94.4505 | -40.9376 | -18.8718 | -10.9439 | -5.5266 | -2.7518 | -1.8269 | -1.9591 | -0.3735 | -1.0342 | -0.3735 | 0.28716 | 4.05 | 0.438 | 4.05 | 0.709 | 4.05 | 0.744
4.05 | 1.2723 | 0.75121 | 4.05 | 0.037 | 0.438 | 0.607 | 0.667 | 0.709 | 0.73 | 0.737 | 0.736 | 0.748 | 0.744 | 0.749 | 0.753 |  | -95.0746 | -41.6937 | -19.1966 | -11.2094 | -5.6184 | -2.8229 | -1.8911 | -2.0242 | -0.42676 | -0.95924 | -0.29364 | 0.23884 | 4.1 | 0.428 | 4.1 | 0.703 | 4.1 | 0.738
4.1 | 1.2881 | 0.74554 | 4.1 | 0.033 | 0.428 | 0.599 | 0.66 | 0.703 | 0.724 | 0.732 | 0.73 | 0.743 | 0.738 | 0.743 | 0.748 |  | -95.5737 | -42.5919 | -19.6554 | -11.4734 | -5.7058 | -2.889 | -1.816 | -2.0843 | -0.34055 | -1.0112 | -0.34055 | 0.3301 | 4.15 | 0.419 | 4.15 | 0.697 | 4.15 | 0.732
4.15 | 1.3038 | 0.73983 | 4.15 | 0.029 | 0.419 | 0.591 | 0.654 | 0.697 | 0.718 | 0.726 | 0.724 | 0.737 | 0.732 | 0.738 | 0.742 |  | -96.0802 | -43.3651 | -20.1164 | -11.6009 | -5.7888 | -2.9503 | -1.8689 | -2.1393 | -0.3821 | -1.0579 | -0.24694 | 0.29373 | 4.2 | 0.41 | 4.2 | 0.69 | 4.2 | 0.727
4.2 | 1.3195 | 0.73407 | 4.2 | 0.025 | 0.41 | 0.583 | 0.647 | 0.69 | 0.712 | 0.72 | 0.718 | 0.731 | 0.727 | 0.732 | 0.736 |  | -96.5943 | -44.1471 | -20.5798 | -11.8613 | -6.0036 | -3.0066 | -1.9168 | -2.1892 | -0.41828 | -0.96319 | -0.28205 | 0.26285 | 4.25 | 0.401 | 4.25 | 0.684 | 4.25 | 0.721
4.25 | 1.3352 | 0.72827 | 4.25 | 0.022 | 0.401 | 0.576 | 0.64 | 0.684 | 0.706 | 0.714 | 0.712 | 0.726 | 0.721 | 0.726 | 0.731 |  | -96.9791 | -44.938 | -20.9085 | -12.1206 | -6.0789 | -3.058 | -1.9595 | -2.2341 | -0.31176 | -0.99832 | -0.31176 | 0.37479 | 4.3 | 0.392 | 4.3 | 0.678 | 4.3 | 0.715
4.3 | 1.3509 | 0.72243 | 4.3 | 0.019 | 0.392 | 0.568 | 0.633 | 0.678 | 0.7 | 0.708 | 0.706 | 0.72 | 0.715 | 0.721 | 0.725 |  | -97.37 | -45.7385 | -21.3762 | -12.3788 | -6.1498 | -3.1045 | -1.9971 | -2.274 | -0.33606 | -1.0282 | -0.19764 | 0.35605 | 4.35 | 0.383 | 4.35 | 0.671 | 4.35 | 0.709
4.35 | 1.3666 | 0.71654 | 4.35 | 0.016 | 0.383 | 0.56 | 0.626 | 0.671 | 0.694 | 0.702 | 0.7 | 0.714 | 0.709 | 0.715 | 0.72 |  | -97.7671 | -46.5489 | -21.847 | -12.6361 | -6.356 | -3.1461 | -2.0296 | -2.3088 | -0.35493 | -1.0527 | -0.21537 | 0.48242 | 4.4 | 0.374 | 4.4 | 0.665 | 4.4 | 0.703
4.4 | 1.3823 | 0.71062 | 4.4 | 0.013 | 0.374 | 0.552 | 0.619 | 0.665 | 0.688 | 0.697 | 0.694 | 0.708 | 0.703 | 0.709 | 0.714 |  | -98.1706 | -47.3697 | -22.3211 | -12.8927 | -6.4194 | -3.1828 | -1.9163 | -2.3385 | -0.36836 | -1.072 | -0.22763 | 0.47598 | 4.45 | 0.366 | 4.45 | 0.659 | 4.45 | 0.697
4.45 | 1.398 | 0.70465 | 4.45 | 0.011 | 0.366 | 0.545 | 0.612 | 0.659 | 0.682 | 0.691 | 0.688 | 0.703 | 0.697 | 0.703 | 0.708 |  | -98.4389 | -48.0595 | -22.6568 | -13.1486 | -6.4786 | -3.2146 | -1.9374 | -2.3631 | -0.23441 | -1.0859 | -0.23441 | 0.47516 | 4.5 | 0.357 | 4.5 | 0.652 | 4.5 | 0.691
4.5 | 1.4137 | 0.69865 | 4.5 | 0.009 | 0.357 | 0.537 | 0.605 | 0.652 | 0.676 | 0.685 | 0.682 | 0.697 | 0.691 | 0.697 | 0.702 |  | -98.7118 | -48.9012 | -23.1371 | -13.404 | -6.6767 | -3.2415 | -1.9533 | -2.3827 | -0.23568 | -1.0945 | -0.23568 | 0.47999 | 4.55 | 0.349 | 4.55 | 0.646 | 4.55 | 0.685
4.55 | 1.4294 | 0.6926 | 4.55 | 0.007 | 0.349 | 0.529 | 0.598 | 0.646 | 0.67 | 0.679 | 0.676 | 0.691 | 0.685 | 0.691 | 0.696 |  | -98.9893 | -49.6104 | -23.6215 | -13.659 | -6.7287 | -3.2635 | -1.964 | -2.3972 | -0.23143 | -1.0977 | -0.23143 | 0.49048 | 4.6 | 0.34 | 4.6 | 0.639 | 4.6 | 0.679
4.6 | 1.4451 | 0.68652 | 4.6 | 0.006 | 0.34 | 0.521 | 0.591 | 0.639 | 0.664 | 0.673 | 0.67 | 0.685 | 0.679 | 0.686 | 0.691 |  | -99.126 | -50.475 | -24.1102 | -13.9138 | -6.9221 | -3.2805 | -1.9696 | -2.4066 | -0.22163 | -1.0956 | -0.075971 | 0.65234 | 4.65 | 0.332 | 4.65 | 0.633 | 4.65 | 0.673
4.65 | 1.4608 | 0.6804 | 4.65 | 0.004 | 0.332 | 0.514 | 0.584 | 0.633 | 0.658 | 0.666 | 0.664 | 0.679 | 0.673 | 0.68 | 0.685 |  | -99.4121 | -51.2054 | -24.4566 | -14.1686 | -6.967 | -3.2927 | -2.1169 | -2.4108 | -0.20627 | -1.0881 | -0.059294 | 0.67556 | 4.7 | 0.324 | 4.7 | 0.626 | 4.7 | 0.667
4.7 | 1.4765 | 0.67425 | 4.7 | 0.003 | 0.324 | 0.506 | 0.577 | 0.626 | 0.652 | 0.66 | 0.658 | 0.673 | 0.667 | 0.674 | 0.679 |  | -99.5551 | -51.9466 | -24.9536 | -14.4234 | -7.156 | -3.2999 | -2.1134 | -2.41 | -0.1853 | -1.0752 | -0.036989 | 0.70458 | 4.75 | 0.315 | 4.75 | 0.62 | 4.75 | 0.661
4.75 | 1.4923 | 0.66806 | 4.75 | 0.002 | 0.315 | 0.498 | 0.57 | 0.62 | 0.645 | 0.654 | 0.651 | 0.667 | 0.661 | 0.668 | 0.673 |  | -99.7006 | -52.8486 | -25.4558 | -14.6784 | -7.194 | -3.4518 | -2.1046 | -2.5537 | -0.15871 | -1.0568 | -0.0090263 | 0.73941 | 4.8 | 0.307 | 4.8 | 0.613 | 4.8 | 0.655
4.8 | 1.508 | 0.66184 | 4.8 | 0.001 | 0.307 | 0.491 | 0.563 | 0.613 | 0.639 | 0.648 | 0.645 | 0.661 | 0.655 | 0.662 | 0.667 |  | -99.8489 | -53.614 | -25.8126 | -14.9337 | -7.379 | -3.4506 | -2.0907 | -2.544 | -0.12647 | -1.033 | 0.024624 | 0.7801 | 4.85 | 0.299 | 4.85 | 0.607 | 4.85 | 0.649
4.85 | 1.5237 | 0.65558 | 4.85 | 0.001 | 0.299 | 0.483 | 0.556 | 0.607 | 0.633 | 0.642 | 0.639 | 0.655 | 0.649 | 0.656 | 0.661 |  | -99.8475 | -54.3916 | -26.3248 | -15.1897 | -7.4103 | -3.4443 | -2.0715 | -2.5291 | -0.088541 | -1.0038 | 0.063995 | 0.82668 | 4.9 | 0.292 | 4.9 | 0.6 | 4.9 | 0.643
4.9 | 1.5394 | 0.64929 | 4.9 | 0 | 0.292 | 0.475 | 0.549 | 0.6 | 0.627 | 0.636 | 0.633 | 0.649 | 0.643 | 0.649 | 0.655 |  | -100 | -55.0279 | -26.8433 | -15.4463 | -7.5916 | -3.4332 | -2.0471 | -2.5091 | -0.04489 | -0.96897 | -0.04489 | 0.87919 | 4.95 | 0.284 | 4.95 | 0.594 | 4.95 | 0.636
4.95 | 1.5551 | 0.64297 | 4.95 | 0 | 0.284 | 0.468 | 0.542 | 0.594 | 0.62 | 0.63 | 0.626 | 0.643 | 0.636 | 0.643 | 0.649 |  | -100 | -55.83 | -27.2129 | -15.7038 | -7.6164 | -3.5726 | -2.0173 | -2.6395 | 0.0045191 | -1.0842 | 0.0045191 | 0.93769 | 5 | 0.276 | 5 | 0.587 | 5 | 0.63
5 | 1.5708 | 0.63662 | 5 | 0 | 0.276 | 0.46 | 0.535 | 0.587 | 0.614 | 0.623 | 0.62 | 0.637 | 0.63 | 0.637 | 0.643 |  | -100 | -56.646 | -27.7434 | -15.9624 | -7.7943 | -3.5531 | -2.1394 | -2.6106 | 0.059726 | -1.0398 | 0.059726 | 1.0022 | 5.05 | 0.269 | 5.05 | 0.581 | 5.05 | 0.624
5.05 | 1.5865 | 0.63024 | 5.05 | 0 | 0.269 | 0.453 | 0.528 | 0.581 | 0.608 | 0.617 | 0.614 | 0.631 | 0.624 | 0.631 | 0.637 |  | -100 | -57.3178 | -28.1225 | -16.2222 | -7.8127 | -3.5286 | -2.1006 | -2.5766 | 0.12077 | -0.98992 | 0.12077 | 1.0728 | 5.1 | 0.261 | 5.1 | 0.574 | 5.1 | 0.618
5.1 | 1.6022 | 0.62383 | 5.1 | 0 | 0.261 | 0.445 | 0.521 | 0.574 | 0.602 | 0.611 | 0.608 | 0.624 | 0.618 | 0.625 | 0.63 |  | -100 | -58.1616 | -28.6664 | -16.4835 | -7.9876 | -3.4992 | -2.0565 | -2.5374 | 0.027402 | -0.9344 | 0.1877 | 0.9892 | 5.15 | 0.254 | 5.15 | 0.568 | 5.15 | 0.611
5.15 | 1.6179 | 0.61739 | 5.15 | 0.001 | 0.254 | 0.438 | 0.514 | 0.568 | 0.595 | 0.605 | 0.601 | 0.618 | 0.611 | 0.619 | 0.624 |  | -99.838 | -58.8592 | -29.0563 | -16.7465 | -8 | -3.6268 | -2.007 | -2.6549 | 0.098591 | -1.0352 | 0.26056 | 1.0704 | 5.2 | 0.247 | 5.2 | 0.561 | 5.2 | 0.605
5.2 | 1.6336 | 0.61093 | 5.2 | 0.001 | 0.247 | 0.431 | 0.507 | 0.561 | 0.589 | 0.599 | 0.595 | 0.612 | 0.605 | 0.613 | 0.618 |  | -99.8363 | -59.5696 | -29.4514 | -17.0113 | -8.1723 | -3.5891 | -1.9522 | -2.6069 | 0.17572 | -0.97008 | 0.3394 | 1.1578 | 5.25 | 0.24 | 5.25 | 0.555 | 5.25 | 0.599
5.25 | 1.6493 | 0.60444 | 5.25 | 0.002 | 0.24 | 0.423 | 0.5 | 0.555 | 0.583 | 0.592 | 0.588 | 0.606 | 0.599 | 0.606 | 0.612 |  | -99.6691 | -60.2935 | -30.0173 | -17.2782 | -8.1788 | -3.5464 | -2.0574 | -2.7192 | 0.25883 | -0.89927 | 0.25883 | 1.2515 | 5.3 | 0.233 | 5.3 | 0.548 | 5.3 | 0.592
5.3 | 1.665 | 0.59792 | 5.3 | 0.003 | 0.233 | 0.416 | 0.493 | 0.548 | 0.576 | 0.586 | 0.582 | 0.6 | 0.592 | 0.6 | 0.606 |  | -99.4983 | -61.0315 | -30.4254 | -17.5474 | -8.3488 | -3.6659 | -1.9935 | -2.6624 | 0.34799 | -0.98998 | 0.34799 | 1.3515 | 5.35 | 0.226 | 5.35 | 0.541 | 5.35 | 0.586
5.35 | 1.6808 | 0.59138 | 5.35 | 0.004 | 0.226 | 0.409 | 0.486 | 0.541 | 0.57 | 0.58 | 0.576 | 0.594 | 0.586 | 0.594 | 0.6 |  | -99.3236 | -61.7842 | -30.8396 | -17.8192 | -8.5189 | -3.6151 | -1.9241 | -2.6005 | 0.44325 | -0.90952 | 0.44325 | 1.4578 | 5.4 | 0.219 | 5.4 | 0.535 | 5.4 | 0.58
5.4 | 1.6965 | 0.58481 | 5.4 | 0.006 | 0.219 | 0.401 | 0.479 | 0.535 | 0.564 | 0.573 | 0.569 | 0.587 | 0.58 | 0.588 | 0.593 |  | -98.974 | -62.5522 | -31.4313 | -18.0937 | -8.518 | -3.5592 | -2.0202 | -2.7042 | 0.37368 | -0.82328 | 0.54467 | 1.3996 | 5.45 | 0.213 | 5.45 | 0.528 | 5.45 | 0.573
5.45 | 1.7122 | 0.57823 | 5.45 | 0.007 | 0.213 | 0.394 | 0.472 | 0.528 | 0.557 | 0.567 | 0.563 | 0.581 | 0.573 | 0.581 | 0.587 |  | -98.7894 | -63.1633 | -31.8608 | -18.3713 | -8.6866 | -3.6712 | -1.9418 | -2.6336 | 0.47938 | -0.90416 | 0.47938 | 1.517 | 5.5 | 0.206 | 5.5 | 0.522 | 5.5 | 0.567
5.5 | 1.7279 | 0.57162 | 5.5 | 0.009 | 0.206 | 0.387 | 0.465 | 0.522 | 0.551 | 0.561 | 0.556 | 0.575 | 0.567 | 0.575 | 0.581 |  | -98.4255 | -63.9621 | -32.2977 | -18.6522 | -8.6806 | -3.6073 | -1.8579 | -2.7326 | 0.59131 | -0.80822 | 0.59131 | 1.641 | 5.55 | 0.2 | 5.55 | 0.515 | 5.55 | 0.561
5.55 | 1.7436 | 0.56499 | 5.55 | 0.011 | 0.2 | 0.38 | 0.458 | 0.515 | 0.545 | 0.555 | 0.55 | 0.569 | 0.561 | 0.569 | 0.575 |  | -98.0531 | -64.6012 | -32.7423 | -18.9368 | -8.8481 | -3.5383 | -1.7684 | -2.6533 | 0.70956 | -0.70639 | 0.70956 | 1.7715 | 5.6 | 0.194 | 5.6 | 0.509 | 5.6 | 0.554
5.6 | 1.7593 | 0.55834 | 5.6 | 0.013 | 0.194 | 0.373 | 0.452 | 0.509 | 0.538 | 0.548 | 0.544 | 0.562 | 0.554 | 0.563 | 0.568 |  | -97.6717 | -65.2543 | -33.1951 | -19.0461 | -8.8373 | -3.6434 | -1.8523 | -2.5687 | 0.65508 | -0.77773 | 0.83418 | 1.7297 | 5.65 | 0.188 | 5.65 | 0.502 | 5.65 | 0.548
5.65 | 1.775 | 0.55167 | 5.65 | 0.016 | 0.188 | 0.366 | 0.445 | 0.502 | 0.532 | 0.542 | 0.537 | 0.556 | 0.548 | 0.556 | 0.562 |  | -97.0997 | -65.922 | -33.6566 | -19.3365 | -9.0044 | -3.5664 | -1.7537 | -2.6601 | 0.784 | -0.66613 | 0.784 | 1.8716 | 5.7 | 0.182 | 5.7 | 0.496 | 5.7 | 0.541
5.7 | 1.7907 | 0.54499 | 5.7 | 0.019 | 0.182 | 0.359 | 0.438 | 0.496 | 0.526 | 0.536 | 0.531 | 0.55 | 0.541 | 0.55 | 0.556 |  | -96.5137 | -66.6049 | -34.1272 | -19.6315 | -8.9891 | -3.4844 | -1.6495 | -2.5669 | 0.9194 | -0.73201 | 0.9194 | 2.0203 | 5.75 | 0.176 | 5.75 | 0.489 | 5.75 | 0.535
5.75 | 1.8064 | 0.53829 | 5.75 | 0.022 | 0.176 | 0.352 | 0.431 | 0.489 | 0.519 | 0.529 | 0.524 | 0.544 | 0.535 | 0.544 | 0.55 |  | -95.913 | -67.3037 | -34.6074 | -19.9312 | -9.1562 | -3.583 | -1.7253 | -2.6541 | 1.0614 | -0.61062 | 1.0614 | 2.176 | 5.8 | 0.17 | 5.8 | 0.483 | 5.8 | 0.529
5.8 | 1.8221 | 0.53157 | 5.8 | 0.025 | 0.17 | 0.345 | 0.425 | 0.483 | 0.513 | 0.523 | 0.518 | 0.537 | 0.529 | 0.537 | 0.543 |  | -95.2969 | -68.0192 | -35.0977 | -20.0479 | -9.1368 | -3.4931 | -1.6119 | -2.5525 | 1.0218 | -0.48315 | 1.0218 | 2.1506 | 5.85 | 0.164 | 5.85 | 0.476 | 5.85 | 0.522
5.85 | 1.8378 | 0.52483 | 5.85 | 0.029 | 0.164 | 0.338 | 0.418 | 0.476 | 0.507 | 0.517 | 0.511 | 0.531 | 0.522 | 0.531 | 0.537 |  | -94.4744 | -68.7521 | -35.5987 | -20.3558 | -9.3047 | -3.3981 | -1.4927 | -2.636 | 1.1748 | -0.54007 | 1.1748 | 2.318 | 5.9 | 0.159 | 5.9 | 0.47 | 5.9 | 0.516
5.9 | 1.8535 | 0.51809 | 5.9 | 0.033 | 0.159 | 0.332 | 0.411 | 0.47 | 0.5 | 0.51 | 0.505 | 0.525 | 0.516 | 0.525 | 0.531 |  | -93.6304 | -69.3101 | -35.918 | -20.6696 | -9.2815 | -3.491 | -1.5608 | -2.5259 | 1.3344 | -0.40271 | 1.3344 | 2.4926 | 5.95 | 0.154 | 5.95 | 0.464 | 5.95 | 0.509
5.95 | 1.8692 | 0.51132 | 5.95 | 0.037 | 0.154 | 0.325 | 0.405 | 0.464 | 0.494 | 0.504 | 0.499 | 0.518 | 0.509 | 0.518 | 0.524 |  | -92.7639 | -69.8822 | -36.4396 | -20.794 | -9.2554 | -3.3883 | -1.4325 | -2.4104 | 1.3054 | -0.4547 | 1.3054 | 2.4789 | 6 | 0.148 | 6 | 0.457 | 6 | 0.503
6 | 1.885 | 0.50455 | 6 | 0.042 | 0.148 | 0.318 | 0.398 | 0.457 | 0.488 | 0.498 | 0.492 | 0.512 | 0.503 | 0.512 | 0.518 |  | -91.6758 | -70.667 | -36.9737 | -21.118 | -9.4244 | -3.2804 | -1.2984 | -2.4876 | 1.4763 | -0.30743 | 1.4763 | 2.6655 | 6.05 | 0.143 | 6.05 | 0.451 | 6.05 | 0.497
6.05 | 1.9007 | 0.49777 | 6.05 | 0.047 | 0.143 | 0.312 | 0.392 | 0.451 | 0.481 | 0.492 | 0.486 | 0.506 | 0.497 | 0.506 | 0.512 |  | -90.5578 | -71.2716 | -37.3199 | -21.2481 | -9.3951 | -3.3682 | -1.1583 | -2.3637 | 1.6542 | -0.15385 | 1.6542 | 2.8596 | 6.1 | 0.138 | 6.1 | 0.444 | 6.1 | 0.49
6.1 | 1.9164 | 0.49097 | 6.1 | 0.052 | 0.138 | 0.305 | 0.385 | 0.444 | 0.475 | 0.485 | 0.479 | 0.5 | 0.49 | 0.499 | 0.505 |  | -89.4087 | -71.8924 | -37.8781 | -21.5838 | -9.5668 | -3.2527 | -1.2159 | -2.438 | 1.8392 | -0.19755 | 1.6356 | 2.8576 | 6.15 | 0.133 | 6.15 | 0.438 | 6.15 | 0.484
6.15 | 1.9321 | 0.48416 | 6.15 | 0.058 | 0.133 | 0.299 | 0.379 | 0.438 | 0.469 | 0.479 | 0.473 | 0.493 | 0.484 | 0.493 | 0.499 |  | -88.0206 | -72.53 | -38.2441 | -21.7208 | -9.5349 | -3.1321 | -1.0667 | -2.3059 | 1.8249 | -0.033948 | 1.8249 | 3.0642 | 6.2 | 0.129 | 6.2 | 0.432 | 6.2 | 0.477
6.2 | 1.9478 | 0.47735 | 6.2 | 0.064 | 0.129 | 0.292 | 0.372 | 0.432 | 0.463 | 0.473 | 0.466 | 0.487 | 0.477 | 0.487 | 0.493 |  | -86.5926 | -72.9758 | -38.829 | -22.0698 | -9.5004 | -3.0062 | -0.9113 | -2.3777 | 2.0216 | -0.073337 | 2.0216 | 3.2785 | 6.25 | 0.124 | 6.25 | 0.426 | 6.25 | 0.471
6.25 | 1.9635 | 0.47053 | 6.25 | 0.071 | 0.124 | 0.286 | 0.366 | 0.426 | 0.456 | 0.467 | 0.46 | 0.481 | 0.471 | 0.48 | 0.486 |  | -84.9106 | -73.6466 | -39.2172 | -22.215 | -9.4634 | -3.0876 | -0.74979 | -2.2375 | 2.2256 | 0.10032 | 2.0131 | 3.2882 | 6.3 | 0.119 | 6.3 | 0.419 | 6.3 | 0.465
6.3 | 1.9792 | 0.4637 | 6.3 | 0.077 | 0.119 | 0.28 | 0.359 | 0.419 | 0.45 | 0.46 | 0.454 | 0.475 | 0.465 | 0.474 | 0.48 |  | -83.3944 | -74.3368 | -39.616 | -22.5791 | -9.6397 | -2.9543 | -0.79772 | -2.0917 | 2.4371 | 0.28057 | 2.2215 | 3.5154 | 6.35 | 0.115 | 6.35 | 0.413 | 6.35 | 0.458
6.35 | 1.9949 | 0.45686 | 6.35 | 0.085 | 0.115 | 0.274 | 0.353 | 0.413 | 0.444 | 0.454 | 0.447 | 0.468 | 0.458 | 0.468 | 0.474 |  | -81.3949 | -74.8284 | -40.0259 | -22.7341 | -9.6011 | -2.8157 | -0.62689 | -2.1591 | 2.4375 | 0.24864 | 2.4375 | 3.7508 | 6.4 | 0.111 | 6.4 | 0.407 | 6.4 | 0.452
6.4 | 2.0106 | 0.45002 | 6.4 | 0.092 | 0.111 | 0.268 | 0.347 | 0.407 | 0.438 | 0.448 | 0.441 | 0.462 | 0.452 | 0.461 | 0.468 |  | -79.5566 | -75.3347 | -40.4476 | -22.893 | -9.5604 | -2.6719 | -0.44977 | -2.0052 | 2.6612 | 0.43908 | 2.439 | 3.9944 | 6.45 | 0.106 | 6.45 | 0.401 | 6.45 | 0.445
6.45 | 2.0263 | 0.44318 | 6.45 | 0.101 | 0.106 | 0.262 | 0.341 | 0.401 | 0.432 | 0.442 | 0.435 | 0.456 | 0.445 | 0.455 | 0.461 |  | -77.2102 | -76.0819 | -40.8818 | -23.0561 | -9.5176 | -2.5227 | -0.26624 | -1.8457 | 2.8927 | 0.41068 | 2.6671 | 4.021 | 6.5 | 0.102 | 6.5 | 0.394 | 6.5 | 0.439
6.5 | 2.042 | 0.43633 | 6.5 | 0.109 | 0.102 | 0.256 | 0.335 | 0.394 | 0.426 | 0.436 | 0.428 | 0.45 | 0.439 | 0.449 | 0.455 |  | -75.0191 | -76.6233 | -41.3292 | -23.2237 | -9.7019 | -2.3681 | -0.076225 | -1.9097 | 3.1323 | 0.61132 | 2.9032 | 4.2783 | 6.55 | 0.098 | 6.55 | 0.388 | 6.55 | 0.433
6.55 | 2.0577 | 0.42948 | 6.55 | 0.118 | 0.098 | 0.25 | 0.328 | 0.388 | 0.419 | 0.429 | 0.422 | 0.443 | 0.433 | 0.443 | 0.449 |  | -72.5251 | -77.1819 | -41.7905 | -23.6291 | -9.6588 | -2.4408 | -0.11245 | -1.7423 | 3.1473 | 0.8189 | 3.1473 | 4.5443 | 6.6 | 0.094 | 6.6 | 0.382 | 6.6 | 0.426
6.6 | 2.0735 | 0.42263 | 6.6 | 0.127 | 0.094 | 0.244 | 0.322 | 0.382 | 0.413 | 0.423 | 0.416 | 0.437 | 0.426 | 0.436 | 0.442 |  | -69.9502 | -77.7584 | -42.2665 | -23.8108 | -9.614 | -2.279 | 0.087088 | -1.5692 | 3.3997 | 0.79693 | 3.1631 | 4.5827 | 6.65 | 0.091 | 6.65 | 0.376 | 6.65 | 0.42
6.65 | 2.0892 | 0.41578 | 6.65 | 0.137 | 0.091 | 0.238 | 0.316 | 0.376 | 0.407 | 0.417 | 0.409 | 0.431 | 0.42 | 0.43 | 0.436 |  | -67.0499 | -78.1134 | -42.7583 | -23.9984 | -9.5677 | -2.1118 | 0.29332 | -1.6308 | 3.6605 | 1.0149 | 3.42 | 4.863 | 6.7 | 0.087 | 6.7 | 0.37 | 6.7 | 0.414
6.7 | 2.1049 | 0.40893 | 6.7 | 0.148 | 0.087 | 0.233 | 0.31 | 0.37 | 0.401 | 0.411 | 0.403 | 0.425 | 0.414 | 0.424 | 0.43 |  | -63.8079 | -78.7249 | -43.0219 | -24.1923 | -9.5198 | -1.9391 | 0.50635 | -1.45 | 3.9299 | 1.24 | 3.6854 | 5.1526 | 6.75 | 0.084 | 6.75 | 0.364 | 6.75 | 0.407
6.75 | 2.1206 | 0.40208 | 6.75 | 0.158 | 0.084 | 0.227 | 0.304 | 0.364 | 0.395 | 0.405 | 0.397 | 0.419 | 0.407 | 0.418 | 0.424 |  | -60.7043 | -79.1086 | -43.5435 | -24.3931 | -9.4707 | -1.7608 | 0.7263 | -1.2634 | 4.2082 | 1.2237 | 3.9595 | 5.4517 | 6.8 | 0.08 | 6.8 | 0.358 | 6.8 | 0.401
6.8 | 2.1363 | 0.39523 | 6.8 | 0.169 | 0.08 | 0.221 | 0.299 | 0.358 | 0.389 | 0.399 | 0.391 | 0.412 | 0.401 | 0.411 | 0.417 |  | -57.2403 | -79.7587 | -44.0835 | -24.3483 | -9.4203 | -1.5769 | 0.9533 | -1.0708 | 4.2425 | 1.4593 | 3.9895 | 5.5076 | 6.85 | 0.077 | 6.85 | 0.352 | 6.85 | 0.395
6.85 | 2.152 | 0.38839 | 6.85 | 0.181 | 0.077 | 0.216 | 0.293 | 0.352 | 0.383 | 0.393 | 0.384 | 0.406 | 0.395 | 0.405 | 0.411 |  | -53.3971 | -80.1745 | -44.3855 | -24.56 | -9.369 | -1.3873 | 1.1875 | -1.1298 | 4.5347 | 1.7024 | 4.2772 | 5.822 | 6.9 | 0.074 | 6.9 | 0.346 | 6.9 | 0.389
6.9 | 2.1677 | 0.38155 | 6.9 | 0.193 | 0.074 | 0.211 | 0.287 | 0.346 | 0.377 | 0.387 | 0.378 | 0.4 | 0.389 | 0.399 | 0.405 |  | -49.4165 | -80.6053 | -44.6989 | -24.7801 | -9.3167 | -1.1919 | 1.429 | -0.92982 | 4.8362 | 1.9532 | 4.5741 | 6.1466 | 6.95 | 0.07 | 6.95 | 0.34 | 6.95 | 0.382
6.95 | 2.1834 | 0.37471 | 6.95 | 0.205 | 0.07 | 0.205 | 0.281 | 0.34 | 0.371 | 0.381 | 0.372 | 0.394 | 0.382 | 0.393 | 0.399 |  | -45.2914 | -81.319 | -45.2914 | -25.0092 | -9.2638 | -0.99074 | 1.678 | -0.72387 | 5.1473 | 1.9448 | 4.8804 | 6.4817 | 7 | 0.067 | 7 | 0.335 | 7 | 0.376
7 | 2.1991 | 0.36788 | 7 | 0.218 | 0.067 | 0.2 | 0.276 | 0.335 | 0.365 | 0.375 | 0.366 | 0.388 | 0.376 | 0.387 | 0.392 |  | -40.742 | -81.7877 | -45.6349 | -24.9761 | -8.9384 | -0.78368 | 1.9346 | -0.51185 | 5.4683 | 2.2064 | 5.1965 | 6.5556 | 7.05 | 0.064 | 7.05 | 0.329 | 7.05 | 0.37
7.05 | 2.2148 | 0.36106 | 7.05 | 0.231 | 0.064 | 0.195 | 0.27 | 0.329 | 0.359 | 0.369 | 0.36 | 0.382 | 0.37 | 0.381 | 0.386 |  | -36.0218 | -82.2744 | -45.9924 | -25.2203 | -8.8795 | -0.57063 | 2.199 | -0.29367 | 5.7995 | 2.4759 | 5.5225 | 6.9073 | 7.1 | 0.062 | 7.1 | 0.323 | 7.1 | 0.364
7.1 | 2.2305 | 0.35425 | 7.1 | 0.244 | 0.062 | 0.19 | 0.265 | 0.323 | 0.354 | 0.363 | 0.354 | 0.376 | 0.364 | 0.374 | 0.38 |  | -31.1212 | -82.498 | -46.3648 | -25.1931 | -8.8202 | -0.069241 | 2.4714 | -0.069241 | 6.1411 | 2.7537 | 5.5766 | 7.2703 | 7.15 | 0.059 | 7.15 | 0.317 | 7.15 | 0.358
7.15 | 2.2462 | 0.34744 | 7.15 | 0.258 | 0.059 | 0.185 | 0.259 | 0.317 | 0.348 | 0.357 | 0.348 | 0.37 | 0.358 | 0.368 | 0.374 |  | -25.7423 | -83.0186 | -46.7532 | -25.4545 | -8.7609 | 0.16154 | 2.7519 | 0.16154 | 6.4936 | 3.0397 | 5.9179 | 7.6449 | 7.2 | 0.056 | 7.2 | 0.312 | 7.2 | 0.352
7.2 | 2.2619 | 0.34064 | 7.2 | 0.272 | 0.056 | 0.18 | 0.254 | 0.312 | 0.342 | 0.351 | 0.342 | 0.364 | 0.352 | 0.362 | 0.368 |  | -20.1507 | -83.5604 | -47.1585 | -25.4348 | -8.4081 | 0.39876 | 3.0408 | 0.39876 | 6.8572 | 3.3344 | 6.27 | 8.0314 | 7.25 | 0.054 | 7.25 | 0.306 | 7.25 | 0.345
7.25 | 2.2777 | 0.33385 | 7.25 | 0.287 | 0.054 | 0.175 | 0.248 | 0.306 | 0.336 | 0.345 | 0.336 | 0.358 | 0.345 | 0.356 | 0.362 |  | -14.0345 | -83.8253 | -47.582 | -25.7162 | -8.3434 | 0.64255 | 3.3383 | 0.64255 | 7.2322 | 3.3383 | 6.6332 | 8.4304 | 7.3 | 0.051 | 7.3 | 0.3 | 7.3 | 0.339
7.3 | 2.2934 | 0.32708 | 7.3 | 0.302 | 0.051 | 0.17 | 0.243 | 0.3 | 0.331 | 0.34 | 0.33 | 0.352 | 0.339 | 0.35 | 0.356 |  | -7.6676 | -84.4074 | -48.0248 | -25.7061 | -8.2791 | 1.1988 | 3.9504 | 0.89301 | 7.6192 | 3.6446 | 7.0077 | 8.8422 | 7.35 | 0.049 | 7.35 | 0.295 | 7.35 | 0.333
7.35 | 2.3091 | 0.32032 | 7.35 | 0.317 | 0.049 | 0.166 | 0.238 | 0.295 | 0.325 | 0.334 | 0.324 | 0.346 | 0.333 | 0.344 | 0.35 |  | -1.0351 | -84.7026 | -48.1761 | -25.6983 | -7.9033 | 1.4625 | 4.2722 | 1.1503 | 8.0185 | 3.96 | 7.3941 | 9.2673 | 7.4 | 0.046 | 7.4 | 0.289 | 7.4 | 0.327
7.4 | 2.3248 | 0.31356 | 7.4 | 0.332 | 0.046 | 0.161 | 0.232 | 0.289 | 0.319 | 0.328 | 0.318 | 0.34 | 0.327 | 0.338 | 0.344 |  | 5.8792 | -85.33 | -48.6549 | -26.0121 | -7.834 | 1.7334 | 4.6036 | 1.4145 | 8.4306 | 4.2847 | 7.7927 | 9.7062 | 7.45 | 0.044 | 7.45 | 0.284 | 7.45 | 0.321
7.45 | 2.3405 | 0.30683 | 7.45 | 0.348 | 0.044 | 0.156 | 0.227 | 0.284 | 0.314 | 0.322 | 0.312 | 0.334 | 0.321 | 0.332 | 0.337 |  | 13.4187 | -85.6597 | -49.1571 | -26.0171 | -7.4399 | 2.3375 | 4.9449 | 1.6857 | 8.8559 | 4.6189 | 8.204 | 9.8336 | 7.5 | 0.042 | 7.5 | 0.279 | 7.5 | 0.315
7.5 | 2.3562 | 0.30011 | 7.5 | 0.364 | 0.042 | 0.152 | 0.222 | 0.279 | 0.308 | 0.317 | 0.306 | 0.328 | 0.315 | 0.326 | 0.331 |  | 21.2907 | -86.0049 | -49.3511 | -26.026 | -7.0327 | 2.6306 | 5.6295 | 1.9642 | 9.2949 | 4.9631 | 8.6285 | 10.2946 | 7.55 | 0.04 | 7.55 | 0.273 | 7.55 | 0.309
7.55 | 2.3719 | 0.2934 | 7.55 | 0.38 | 0.04 | 0.148 | 0.217 | 0.273 | 0.302 | 0.311 | 0.3 | 0.322 | 0.309 | 0.32 | 0.325 |  | 29.5166 | -86.3667 | -49.5567 | -26.0392 | -6.9525 | 2.9317 | 5.9992 | 2.25 | 9.7483 | 5.3175 | 9.0667 | 10.7708 | 7.6 | 0.038 | 7.6 | 0.268 | 7.6 | 0.303
7.6 | 2.3876 | 0.28671 | 7.6 | 0.397 | 0.038 | 0.143 | 0.212 | 0.268 | 0.297 | 0.305 | 0.294 | 0.317 | 0.303 | 0.315 | 0.32 |  | 38.4684 | -86.7461 | -50.1235 | -26.0572 | -6.5251 | 3.5897 | 6.38 | 2.5433 | 10.5654 | 5.6824 | 9.8679 | 11.6118 | 7.65 | 0.036 | 7.65 | 0.263 | 7.65 | 0.297
7.65 | 2.4033 | 0.28003 | 7.65 | 0.414 | 0.036 | 0.139 | 0.207 | 0.263 | 0.292 | 0.3 | 0.289 | 0.311 | 0.297 | 0.309 | 0.314 |  | 47.8388 | -87.1444 | -50.3633 | -26.0806 | -6.0831 | 4.2728 | 7.1296 | 3.2015 | 11.0577 | 6.0583 | 10.3435 | 12.129 | 7.7 | 0.034 | 7.7 | 0.257 | 7.7 | 0.292
7.7 | 2.419 | 0.27338 | 7.7 | 0.431 | 0.034 | 0.135 | 0.203 | 0.257 | 0.286 | 0.294 | 0.283 | 0.305 | 0.292 | 0.303 | 0.308 |  | 57.6564 | -87.5631 | -50.6181 | -25.7442 | -5.9914 | 4.6165 | 7.5429 | 3.5192 | 11.5666 | 6.8113 | 10.835 | 12.664 | 7.75 | 0.032 | 7.75 | 0.252 | 7.75 | 0.286
7.75 | 2.4347 | 0.26674 | 7.75 | 0.448 | 0.032 | 0.131 | 0.198 | 0.252 | 0.281 | 0.289 | 0.277 | 0.299 | 0.286 | 0.297 | 0.302 |  | 67.952 | -88.0034 | -50.889 | -25.7712 | -5.527 | 5.3449 | 8.344 | 3.8453 | 12.093 | 7.2194 | 11.3432 | 13.2176 | 7.8 | 0.031 | 7.8 | 0.247 | 7.8 | 0.28
7.8 | 2.4504 | 0.26013 | 7.8 | 0.465 | 0.031 | 0.127 | 0.193 | 0.247 | 0.275 | 0.283 | 0.271 | 0.294 | 0.28 | 0.291 | 0.296 |  | 78.7595 | -88.0827 | -51.1775 | -25.8052 | -5.0461 | 5.718 | 8.7934 | 4.1802 | 13.0221 | 7.6401 | 11.8688 | 13.791 | 7.85 | 0.029 | 7.85 | 0.242 | 7.85 | 0.274
7.85 | 2.4662 | 0.25353 | 7.85 | 0.483 | 0.029 | 0.123 | 0.188 | 0.242 | 0.27 | 0.278 | 0.266 | 0.288 | 0.274 | 0.286 | 0.29 |  | 90.5101 | -88.5615 | -51.485 | -25.847 | -4.5477 | 6.4963 | 9.6518 | 4.9186 | 13.5961 | 8.0741 | 12.8072 | 14.385 | 7.9 | 0.028 | 7.9 | 0.237 | 7.9 | 0.268
7.9 | 2.4819 | 0.24695 | 7.9 | 0.5 | 0.028 | 0.119 | 0.184 | 0.237 | 0.265 | 0.272 | 0.26 | 0.282 | 0.268 | 0.28 | 0.284 |  | 102.4661 | -88.6619 | -51.8131 | -25.4925 | -4.0311 | 7.307 | 10.1416 | 5.2824 | 14.1909 | 8.5218 | 13.381 | 15.0008 | 7.95 | 0.026 | 7.95 | 0.232 | 7.95 | 0.263
7.95 | 2.4976 | 0.2404 | 7.95 | 0.518 | 0.026 | 0.115 | 0.179 | 0.232 | 0.26 | 0.267 | 0.255 | 0.277 | 0.263 | 0.274 | 0.279 |  | 115.4723 | -89.1848 | -52.1635 | -25.5414 | -3.495 | 8.1521 | 11.0639 | 6.0723 | 15.2236 | 9.4 | 13.9757 | 16.0555 | 8 | 0.025 | 8 | 0.227 | 8 | 0.257
8 | 2.5133 | 0.23387 | 8 | 0.536 | 0.025 | 0.112 | 0.175 | 0.227 | 0.254 | 0.262 | 0.249 | 0.271 | 0.257 | 0.269 | 0.273 |  | 129.1849 | -89.3104 | -52.1106 | -25.1728 | -2.9385 | 8.6063 | 12.0269 | 6.4683 | 15.8752 | 9.889 | 15.02 | 16.7304 | 8.05 | 0.023 | 8.05 | 0.222 | 8.05 | 0.251
8.05 | 2.529 | 0.22737 | 8.05 | 0.553 | 0.023 | 0.108 | 0.17 | 0.222 | 0.249 | 0.256 | 0.244 | 0.266 | 0.251 | 0.263 | 0.267 |  | 143.2199 | -89.8842 | -52.4996 | -25.2308 | -2.3602 | 9.5149 | 12.5937 | 7.3158 | 16.9918 | 10.3946 | 15.6724 | 17.4317 | 8.1 | 0.022 | 8.1 | 0.217 | 8.1 | 0.246
8.1 | 2.5447 | 0.22088 | 8.1 | 0.571 | 0.022 | 0.104 | 0.166 | 0.217 | 0.244 | 0.251 | 0.238 | 0.26 | 0.246 | 0.257 | 0.261 |  | 158.5058 | -90.0401 | -52.9166 | -24.8477 | -1.7587 | 10.4648 | 13.6339 | 7.7485 | 17.7084 | 11.3703 | 16.3502 | 18.1611 | 8.15 | 0.021 | 8.15 | 0.212 | 8.15 | 0.24
8.15 | 2.5604 | 0.21443 | 8.15 | 0.589 | 0.021 | 0.101 | 0.162 | 0.212 | 0.239 | 0.246 | 0.233 | 0.255 | 0.24 | 0.252 | 0.256 |  | 174.6834 | -90.2065 | -52.8981 | -24.4504 | -1.1326 | 11.459 | 14.7234 | 8.6608 | 18.9206 | 11.9253 | 17.5216 | 19.387 | 8.2 | 0.02 | 8.2 | 0.208 | 8.2 | 0.235
8.2 | 2.5761 | 0.208 | 8.2 | 0.606 | 0.02 | 0.098 | 0.158 | 0.208 | 0.234 | 0.241 | 0.227 | 0.249 | 0.235 | 0.246 | 0.25 |  | 191.3479 | -90.3846 | -52.8843 | -24.038 | 0.00060617 | 12.5007 | 15.8661 | 9.1353 | 19.7123 | 12.9815 | 18.2699 | 20.193 | 8.25 | 0.018 | 8.25 | 0.203 | 8.25 | 0.229
8.25 | 2.5918 | 0.2016 | 8.25 | 0.624 | 0.018 | 0.094 | 0.153 | 0.203 | 0.229 | 0.236 | 0.222 | 0.244 | 0.229 | 0.241 | 0.245 |  | 209.5304 | -91.0712 | -53.372 | -24.1055 | 0.69659 | 13.5937 | 17.066 | 10.1214 | 21.0343 | 13.5937 | 19.5462 | 21.5304 | 8.3 | 0.017 | 8.3 | 0.198 | 8.3 | 0.224
8.3 | 2.6075 | 0.19522 | 8.3 | 0.641 | 0.017 | 0.091 | 0.149 | 0.198 | 0.224 | 0.231 | 0.217 | 0.239 | 0.224 | 0.235 | 0.239 |  | 228.3469 | -91.2919 | -53.386 | -23.676 | 1.4238 | 14.7421 | 18.3278 | 11.1564 | 22.4257 | 14.7421 | 20.3768 | 22.4257 | 8.35 | 0.016 | 8.35 | 0.194 | 8.35 | 0.218
8.35 | 2.6232 | 0.18887 | 8.35 | 0.659 | 0.016 | 0.088 | 0.145 | 0.194 | 0.219 | 0.226 | 0.212 | 0.233 | 0.218 | 0.23 | 0.233 |  | 248.9107 | -91.5287 | -53.408 | -23.229 | 2.7142 | 15.9506 | 19.6568 | 12.2444 | 23.363 | 15.4212 | 21.7746 | 23.363 | 8.4 | 0.015 | 8.4 | 0.189 | 8.4 | 0.213
8.4 | 2.6389 | 0.18256 | 8.4 | 0.676 | 0.015 | 0.085 | 0.141 | 0.189 | 0.214 | 0.221 | 0.206 | 0.228 | 0.213 | 0.225 | 0.228 |  | 270.2975 | -91.7833 | -53.4389 | -22.7634 | 3.5299 | 17.2244 | 21.0588 | 12.8421 | 24.8933 | 16.6766 | 23.2499 | 24.8933 | 8.45 | 0.014 | 8.45 | 0.185 | 8.45 | 0.207
8.45 | 2.6546 | 0.17627 | 8.45 | 0.693 | 0.014 | 0.082 | 0.137 | 0.185 | 0.21 | 0.216 | 0.201 | 0.223 | 0.207 | 0.219 | 0.223 |  | 293.1507 | -92.0576 | -53.48 | -22.2776 | 4.9537 | 19.1366 | 22.5405 | 14.0307 | 26.5117 | 17.4346 | 24.2424 | 26.5117 | 8.5 | 0.013 | 8.5 | 0.18 | 8.5 | 0.202
8.5 | 2.6704 | 0.17001 | 8.5 | 0.71 | 0.013 | 0.079 | 0.134 | 0.18 | 0.205 | 0.211 | 0.196 | 0.218 | 0.202 | 0.214 | 0.217 |  | 317.6191 | -92.3535 | -53.5325 | -21.1817 | 5.8753 | 20.5802 | 24.1093 | 15.2864 | 28.2267 | 18.8156 | 25.8739 | 27.6385 | 8.55 | 0.013 | 8.55 | 0.176 | 8.55 | 0.197
8.55 | 2.6861 | 0.16379 | 8.55 | 0.727 | 0.013 | 0.076 | 0.13 | 0.176 | 0.2 | 0.206 | 0.191 | 0.213 | 0.197 | 0.209 | 0.212 |  | 343.872 | -92.0628 | -53.598 | -20.6281 | 7.4573 | 22.1106 | 25.7739 | 16.6156 | 30.0478 | 20.2789 | 27.6056 | 29.4372 | 8.6 | 0.012 | 8.6 | 0.172 | 8.6 | 0.192
8.6 | 2.7018 | 0.15759 | 8.6 | 0.743 | 0.012 | 0.073 | 0.126 | 0.172 | 0.196 | 0.201 | 0.186 | 0.208 | 0.192 | 0.204 | 0.206 |  | 371.4684 | -92.3854 | -53.6781 | -20.0471 | 9.1421 | 24.3712 | 27.5439 | 18.0257 | 31.9858 | 21.833 | 29.4476 | 30.7167 | 8.65 | 0.011 | 8.65 | 0.167 | 8.65 | 0.186
8.65 | 2.7175 | 0.15143 | 8.65 | 0.759 | 0.011 | 0.07 | 0.123 | 0.167 | 0.191 | 0.196 | 0.181 | 0.202 | 0.186 | 0.199 | 0.201 |  | 401.2135 | -92.736 | -53.7748 | -18.7757 | 10.2802 | 26.1288 | 29.4306 | 19.5252 | 33.3928 | 22.827 | 31.4117 | 32.7324 | 8.7 | 0.01 | 8.7 | 0.163 | 8.7 | 0.181
8.7 | 2.7332 | 0.14531 | 8.7 | 0.775 | 0.01 | 0.068 | 0.119 | 0.163 | 0.186 | 0.191 | 0.176 | 0.197 | 0.181 | 0.193 | 0.196 |  | 433.3577 | -93.118 | -53.2022 | -18.1038 | 12.1772 | 28.0058 | 31.4469 | 21.1238 | 35.5761 | 24.5648 | 32.8233 | 34.8879 | 8.75 | 0.01 | 8.75 | 0.159 | 8.75 | 0.176
8.75 | 2.7489 | 0.13921 | 8.75 | 0.79 | 0.01 | 0.065 | 0.115 | 0.159 | 0.182 | 0.187 | 0.171 | 0.192 | 0.176 | 0.188 | 0.191 |  | 467.4732 | -92.8168 | -53.3092 | -17.3931 | 14.213 | 30.7343 | 34.3259 | 22.8328 | 37.9175 | 26.4244 | 35.0443 | 37.1992 | 8.8 | 0.009 | 8.8 | 0.155 | 8.8 | 0.171
8.8 | 2.7646 | 0.13316 | 8.8 | 0.805 | 0.009 | 0.062 | 0.112 | 0.155 | 0.177 | 0.182 | 0.167 | 0.187 | 0.171 | 0.183 | 0.185 |  | 504.552 | -93.241 | -53.4382 | -15.8884 | 16.4044 | 32.9263 | 36.6813 | 25.4164 | 40.4363 | 28.4204 | 37.4323 | 38.9343 | 8.85 | 0.008 | 8.85 | 0.151 | 8.85 | 0.166
8.85 | 2.7803 | 0.12714 | 8.85 | 0.82 | 0.008 | 0.06 | 0.109 | 0.151 | 0.173 | 0.178 | 0.162 | 0.183 | 0.166 | 0.178 | 0.18 |  | 544.9833 | -93.7075 | -52.8061 | -14.2644 | 18.7713 | 36.0757 | 40.0086 | 27.4235 | 43.9414 | 30.5698 | 40.0086 | 41.5817 | 8.9 | 0.008 | 8.9 | 0.147 | 8.9 | 0.161
8.9 | 2.796 | 0.12115 | 8.9 | 0.834 | 0.008 | 0.058 | 0.105 | 0.147 | 0.169 | 0.173 | 0.157 | 0.178 | 0.161 | 0.173 | 0.175 |  | 588.402 | -93.3966 | -52.1255 | -13.3307 | 21.337 | 39.4963 | 42.798 | 29.5913 | 46.9251 | 32.8929 | 42.798 | 44.4489 | 8.95 | 0.007 | 8.95 | 0.143 | 8.95 | 0.156
8.95 | 2.8117 | 0.1152 | 8.95 | 0.848 | 0.007 | 0.055 | 0.102 | 0.143 | 0.164 | 0.168 | 0.152 | 0.173 | 0.156 | 0.169 | 0.17 |  | 636.096 | -93.9237 | -52.2579 | -11.4602 | 24.1294 | 42.3582 | 45.8303 | 31.9417 | 50.1705 | 35.4139 | 46.6984 | 47.5664 | 9 | 0.007 | 9 | 0.139 | 9 | 0.151
9 | 2.8274 | 0.10929 | 9 | 0.861 | 0.007 | 0.053 | 0.099 | 0.139 | 0.16 | 0.164 | 0.148 | 0.168 | 0.151 | 0.164 | 0.165 |  | 687.7949 | -93.5952 | -51.5062 | -9.4173 | 27.1818 | 46.3963 | 50.0562 | 35.4165 | 53.7161 | 38.1615 | 50.0562 | 50.9711 | 9.05 | 0.006 | 9.05 | 0.135 | 9.05 | 0.147
9.05 | 2.8431 | 0.10342 | 9.05 | 0.874 | 0.006 | 0.051 | 0.096 | 0.135 | 0.156 | 0.16 | 0.143 | 0.163 | 0.147 | 0.159 | 0.16 |  | 745.0901 | -94.1985 | -50.687 | -7.1755 | 30.5345 | 50.8399 | 54.7076 | 38.2699 | 57.6083 | 42.1376 | 53.7406 | 54.7076 | 9.1 | 0.006 | 9.1 | 0.132 | 9.1 | 0.142
9.1 | 2.8588 | 0.097589 | 9.1 | 0.886 | 0.006 | 0.049 | 0.093 | 0.132 | 0.152 | 0.155 | 0.139 | 0.159 | 0.142 | 0.154 | 0.155 |  | 807.8929 | -93.8517 | -49.7892 | -4.702 | 35.2617 | 55.7559 | 58.83 | 42.4347 | 62.9289 | 45.5088 | 57.8053 | 58.83 | 9.15 | 0.005 | 9.15 | 0.128 | 9.15 | 0.137
9.15 | 2.8746 | 0.091796 | 9.15 | 0.898 | 0.005 | 0.046 | 0.09 | 0.128 | 0.147 | 0.151 | 0.134 | 0.154 | 0.137 | 0.149 | 0.15 |  | 878.2554 | -94.5531 | -49.8889 | -1.9566 | 39.4395 | 60.1376 | 64.4951 | 45.9758 | 67.7632 | 49.2439 | 62.3163 | 63.4057 | 9.2 | 0.005 | 9.2 | 0.124 | 9.2 | 0.132
9.2 | 2.8903 | 0.086044 | 9.2 | 0.909 | 0.005 | 0.044 | 0.087 | 0.124 | 0.143 | 0.147 | 0.13 | 0.15 | 0.132 | 0.145 | 0.145 |  | 956.4366 | -94.189 | -48.8634 | 1.1111 | 44.1124 | 66.1941 | 70.8429 | 51.0855 | 74.3295 | 53.4099 | 68.5185 | 68.5185 | 9.25 | 0.004 | 9.25 | 0.121 | 9.25 | 0.128
9.25 | 2.906 | 0.080333 | 9.25 | 0.92 | 0.004 | 0.042 | 0.084 | 0.121 | 0.139 | 0.142 | 0.125 | 0.145 | 0.128 | 0.14 | 0.141 |  | 1045.2339 | -95.0207 | -47.7176 | 4.5648 | 50.6231 | 73.0299 | 76.7644 | 55.6024 | 80.4988 | 59.3369 | 74.2747 | 75.5195 | 9.3 | 0.004 | 9.3 | 0.117 | 9.3 | 0.123
9.3 | 2.9217 | 0.074664 | 9.3 | 0.93 | 0.004 | 0.041 | 0.081 | 0.117 | 0.135 | 0.138 | 0.121 | 0.141 | 0.123 | 0.136 | 0.136 |  | 1145.5868 | -94.6426 | -45.087 | 8.4866 | 56.7029 | 80.811 | 84.829 | 62.0602 | 88.847 | 64.7389 | 82.1503 | 82.1503 | 9.35 | 0.004 | 9.35 | 0.113 | 9.35 | 0.119
9.35 | 2.9374 | 0.069037 | 9.35 | 0.939 | 0.004 | 0.039 | 0.078 | 0.113 | 0.132 | 0.134 | 0.117 | 0.136 | 0.119 | 0.131 | 0.131 |  | 1260.1485 | -94.206 | -43.5082 | 12.9836 | 63.6813 | 91.203 | 94.1 | 69.4754 | 96.997 | 72.3724 | 89.7545 | 89.7545 | 9.4 | 0.003 | 9.4 | 0.11 | 9.4 | 0.114
9.4 | 2.9531 | 0.063452 | 9.4 | 0.948 | 0.003 | 0.037 | 0.075 | 0.11 | 0.128 | 0.13 | 0.113 | 0.132 | 0.114 | 0.127 | 0.126 |  | 1394.0316 | -95.2721 | -41.6886 | 18.1987 | 73.3581 | 101.7258 | 104.8778 | 78.086 | 108.0297 | 79.662 | 100.1498 | 98.5738 | 9.45 | 0.003 | 9.45 | 0.107 | 9.45 | 0.11
9.45 | 2.9688 | 0.057912 | 9.45 | 0.956 | 0.003 | 0.035 | 0.073 | 0.107 | 0.124 | 0.126 | 0.108 | 0.127 | 0.11 | 0.122 | 0.122 |  | 1550.7837 | -94.8197 | -39.5634 | 26.0536 | 84.7635 | 114.1184 | 117.5719 | 86.4902 | 119.2987 | 89.9437 | 110.6649 | 110.6649 | 9.5 | 0.003 | 9.5 | 0.103 | 9.5 | 0.105
9.5 | 2.9845 | 0.052415 | 9.5 | 0.964 | 0.003 | 0.033 | 0.07 | 0.103 | 0.12 | 0.122 | 0.104 | 0.123 | 0.105 | 0.118 | 0.117 |  | 1739.1539 | -94.2765 | -37.0414 | 33.5485 | 96.5071 | 128.9403 | 132.756 | 98.4149 | 134.6638 | 100.3228 | 125.1246 | 123.2168 | 9.55 | 0.002 | 9.55 | 0.1 | 9.55 | 0.101
9.55 | 3.0002 | 0.046964 | 9.55 | 0.97 | 0.002 | 0.032 | 0.067 | 0.1 | 0.116 | 0.118 | 0.1 | 0.119 | 0.101 | 0.114 | 0.113 |  | 1965.4286 | -95.7414 | -31.8621 | 42.6636 | 112.9308 | 146.9997 | 151.2583 | 112.9308 | 153.3876 | 115.0601 | 142.7411 | 140.6118 | 9.6 | 0.002 | 9.6 | 0.097 | 9.6 | 0.097
9.6 | 3.0159 | 0.041557 | 9.6 | 0.977 | 0.002 | 0.03 | 0.065 | 0.097 | 0.113 | 0.114 | 0.096 | 0.115 | 0.097 | 0.109 | 0.108 |  | 2250.9827 | -95.1873 | -27.8102 | 56.4113 | 133.4138 | 171.9151 | 174.3214 | 131.0075 | 176.7277 | 133.4138 | 162.2898 | 159.8834 | 9.65 | 0.002 | 9.65 | 0.094 | 9.65 | 0.092
9.65 | 3.0316 | 0.036196 | 9.65 | 0.982 | 0.002 | 0.029 | 0.063 | 0.094 | 0.109 | 0.11 | 0.092 | 0.11 | 0.092 | 0.105 | 0.104 |  | 2612.9778 | -94.4746 | -19.8815 | 74.0505 | 159.6944 | 201.135 | 203.8977 | 154.169 | 203.8977 | 154.169 | 190.0842 | 187.3215 | 9.7 | 0.001 | 9.7 | 0.091 | 9.7 | 0.088
9.7 | 3.0473 | 0.030882 | 9.7 | 0.987 | 0.001 | 0.027 | 0.06 | 0.091 | 0.106 | 0.107 | 0.088 | 0.106 | 0.088 | 0.101 | 0.099 |  | 3096.0294 | -96.7619 | -12.5706 | 94.2875 | 194.6694 | 243.2413 | 246.4794 | 184.955 | 243.2413 | 184.955 | 227.0506 | 220.5744 | 9.75 | 0.001 | 9.75 | 0.087 | 9.75 | 0.084
9.75 | 3.0631 | 0.025615 | 9.75 | 0.991 | 0.001 | 0.026 | 0.058 | 0.087 | 0.102 | 0.103 | 0.085 | 0.102 | 0.084 | 0.097 | 0.095 |  | 3768.8763 | -96.096 | 1.5043 | 126.4327 | 239.6491 | 298.2093 | 302.1133 | 231.8411 | 298.2093 | 227.937 | 278.6892 | 270.8812 | 9.8 | 0.001 | 9.8 | 0.084 | 9.8 | 0.08
9.8 | 3.0788 | 0.020395 | 9.8 | 0.994 | 0.001 | 0.024 | 0.056 | 0.084 | 0.099 | 0.099 | 0.081 | 0.098 | 0.08 | 0.093 | 0.091 |  | 4773.8062 | -95.0968 | 17.6774 | 174.5806 | 311.8709 | 385.4193 | 385.4193 | 297.1613 | 380.5161 | 292.258 | 356 | 346.1935 | 9.85 | 0.001 | 9.85 | 0.082 | 9.85 | 0.076
9.85 | 3.0945 | 0.015223 | 9.85 | 0.997 | 0.001 | 0.023 | 0.053 | 0.082 | 0.095 | 0.095 | 0.077 | 0.094 | 0.076 | 0.089 | 0.086 |  | 6449.3904 | -93.4309 | 51.0892 | 248.1622 | 438.666 | 524.0643 | 524.0643 | 405.8205 | 517.4952 | 399.2514 | 484.6497 | 464.9424 | 9.9 | 0 | 9.9 | 0.079 | 9.9 | 0.072
9.9 | 3.1102 | 0.010099 | 9.9 | 0.998 | 0 | 0.022 | 0.051 | 0.079 | 0.092 | 0.092 | 0.073 | 0.091 | 0.072 | 0.085 | 0.082 |  | 9781.8254 | -100 | 117.8358 | 404.9831 | 682.2287 | 810.9498 | 810.9498 | 622.8189 | 801.0482 | 612.9173 | 741.6384 | 711.9336 | 9.95 | 0 | 9.95 | 0.076 | 9.95 | 0.068
9.95 | 3.1259 | 0.0050249 | 9.95 | 1 | 0 | 0.021 | 0.049 | 0.076 | 0.088 |  | 0.07 | 0.087 | 0.068 | 0.081 | 0.078 |  | 19800.8184 | -100 | 317.9172 | 875.1401 | 1412.4622 | 1651.272 | -100 | 1293.0573 | 1631.3712 | 1253.2556 | 1511.9663 | 1452.2638